Chroma 1102x LCR Meter

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Chroma 1102x LCR Meter.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

MXI

MXI2

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Chroma 1102x LCR Meter

instrument.  The functions appear in alphabetical order, with a description

of the function and its C syntax, a description of each parameter, and a list

of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                           Function Name:

Initialize With RS232                    Chr1102x_init

Initialize With Options                  Chr1102x_InitWithOptions

Common Command

*CLS                                  Chr1102x_CLS

Set ESE                               Chr1102x_SetESE

Set OPC                               Chr1102x_SetOPC

Set RST                               Chr1102x_SetRST

Set SRE                               Chr1102x_SetSRE

Set Trigger                           Chr1102x_SetTrigger

Set Recall Number                     Chr1102x_SetRecallNumber

Set Save Number                       Chr1102x_SetSaveNumber

Query

Queries ESR                        Chr1102x_Q_ESR

Queries ESE                        Chr1102x_Q_ESE

Queries ID String                  Chr1102x_QueriesID_String

Queries OPC                        Chr1102x_Q_OPC

Queries SRE                        Chr1102x_Q_SRE

Queries TST                        Chr1102x_Q_TST

Queries STB                        Chr1102x_Q_STB

ABORT

Set Abort                             Chr1102x_SetAbort

BINing

Set Binning Mode                      Chr1102x_SetBinningMode

Set Upper Limit of Bin                Chr1102x_SetUpperLimitofBin

Set AUX_HI of Bin                     Chr1102x_SetAUX_HIofBin

Set Low Limit of Bin                  Chr1102x_SetLowLimitofBin

Set AUX_LO of Bin                     Chr1102x_SetAUX_LOofBin

Set Nominal of Bin                    Chr1102x_SetNominalofBin

Set BIN More 9 State                  Chr1102x_SetBINMore9State

Set BIN Number (BIN More 9)           Chr1102x_SetBINNumber_9

Set BIN HI Limit (BIN More 9)         Chr1102x_SetBINHI9

Set BIN Lo Limit (BIN More 9)         Chr1102x_SetBINLo9

Set Binning State                     Chr1102x_SetBinningState

Query

Query Binning Mode                 Chr1102x_QueryBinningMode

Query Upper Linit of Bin           Chr1102x_QueryUpperLinitofBin

Query AUX_HI of Bin                Chr1102x_QueryAUX_HI

Query Low Limit of Bin             Chr1102x_QueryLowLimitBin

Query AUX_LO of Bin                Chr1102x_QueryAUX_LO

Query Nominal of Bin               Chr1102x_QueryNominal

Query BIN More 9 State             Chr1102x_Query_BIN9_State

Query BIN Number (BIN More 9)      Chr1102x_QueryBINNumber_9

Query BIN HI Limit (BIN More 9)    Chr1102x_Q_BINHI9

Query BIN Lo Limit (BIN More 9)    Chr1102x_Q_BINLo9

Query Binning State                Chr1102x_Q_BinningState

Query Binning Result               Chr1102x_GetBinningResult

Calculate

Set Item of Measurement               Chr1102x_SetItemofMeasurement

Set Para 1 of Measurement             Chr1102x_SetPara1ofMeasurement

Set Para 2 of Measurement             Chr1102x_SetPara2ofMeasurement

Set Buzzer Output Define              Chr1102x_SetBuzzerOutputDefine

Set Buzzer Output Action              Chr1102x_SetBuzzerOutputAction

Set PRI_LO_LIMIT of Compare           Chr1102x_SetPRI_LO_LIMITofCompare

Set Compare Limit Action              Chr1102x_SetCompareLimitAction

Set PRI_HI_LIMIT of Compare           Chr1102x_SetPRI_HI_LIMITofCompare

Set Num Represent Method              Chr1102x_SetNumericalRepresentMethod

Set Num Represent Meth Action         Chr1102x_SetNumRepresentMethodAction

Set Vm_Im Display State               Chr1102x_SetVm_ImDisplayState

Query

Query Para 1 of Measurement        Chr1102x_Q_Para1

Query Para 2 of Measurement        Chr1102x_Q_Para2

Query Buzzer Output Define         Chr1102x_Q_Buzzer_Outdef

Query Buzzer Output Action         Chr1102x_Q_BuzzerOutAct

Query PRI_LO_Limit of Compare      Chr1102x_Q_PRI_LO

Query Compare Limit Action         Chr1102x_Q_compareLimitAction

Query PRI_HI_LIMIT of Compare      Chr1102x_Q_PRI_HI_LIMIT

Query Num Represent Method         Chr1102x_Q_NumRepMeth

Query Num Represent Meth Action    Chr1102x_Q_NumRepMetAct

Query Vm_Im Display State          Chr1102x_Q_VIDisplay

Query Bias DEV/PCNT                Chr1102x_QueryBias

Query Calc Path                    Chr1102x_Q_calcpath

CALibration

Set Length of Test Cable              Chr1102x_SetLengthofTestCable

Query

Query Length of Test Cable         Chr1102x_Q_lengthoftest

DATA

Set Nominal Value of Compare          Chr1102x_SetNominalValueofCompare

Query

Query Data Value of Compare        Chr1102x_QueryDataValueofCompare

Display

Set LCD Display                       Chr1102x_SetLCDDisplay

Set LCD Text1 Page                    Chr1102x_SetLCDText1Page

Set LCD Text2 Page                    Chr1102x_SetLCDText2Page

Query

Query LCD Display                  Chr1102x_Q_LCDDisplay

Query LCD Text1 Page               Chr1102x_Q_LCDPage

Query LCD Text2 Page               Chr1102x_Q_LCDText2Page

Dual

Set Calc1 of Dual First Measure       Chr1102x_SetCalc1ofDualFirstMeasure

Set Calc2 of Dual First Measure       Chr1102x_SetCalc2ofDualFirstMeasure

Set Calc2 of Dual 2nd Measure         Chr1102x_SetCalc2ofDual2ndMeasure

Set Dual Auto of First Measure        Chr1102x_SetDualAutoofFirstMeasure

Set Dual First Measure Range          Chr1102x_SetDualFirstMeasureRange

Set Dual Auto of Second Measure       Chr1102x_SetDualAutoofSecondMeasure

Set Dual Second Measure Range         Chr1102x_SetDualSecondMeasureRange

Set Dual First Source Freq.           Chr1102x_SetDualFirstSourceFreq

Set Dual Second Source Freq.          Chr1102x_SetDualSecondSourceFreq

Set Dual Binning Mode                 Chr1102x_SetDualBinningMode

Set Dual First Binning Nominal        Chr1102x_SetDualFirstBinningNominal

Set Dual First Binning Aux_Hi         Chr1102x_SetDualFirstBinningAux_Hi

Set Dual First Binning Aux_Lo         Chr1102x_SetDualFirstBinningAux_Lo

Set Dual Second Binning Aux_Hi        Chr1102x_SetDualSecondBinningAux_Hi

Set Dual Second Binning Aux_Lo        Chr1102x_SetDualSecondBinningAux_Lo

Set First Binning Upper               Chr1102x_SetFirstBinningUpper

Set First Binning Lower               Chr1102x_SetFirstBinningLower

Set Trigger Source of Dual            Chr1102x_SetTriggerSourceofDual

Set Measure Speed of Dual             Chr1102x_SetMeasureSpeedofDual

Query

Calc1 of Dual First Measure        Chr1102x_Q_Calc1DualFirst

Calc2 of Dual First Measure        Chr1102x_Q_Calc2DualFirst

Calc2 of Dual 2nd Measure          Chr1102x_Q_Calc2Dual2nd

Dual Auto of First Measure         Chr1102x_DualAutoFirstMea

Dual First Measure Range           Chr1102x_DualFirstMeaRange

Dual Auto of Second Measure        Chr1102x_DualAutoSecMea

Dual Second Measure Range          Chr1102x_DualSecMeaRange

Dual First Source Freq             Chr1102x_DualFirstSouFreq

Dual Second Source Freq            Chr1102x_DualSecSouFreq

Dual Binning Mode                  Chr1102x_DualBinningMode

Dual First Binning Nominal         Chr1102x_DualFirstBinNom

Dual First Binning Aux_Hi          Chr1102x_DualFriBinAux

Dual First Binning Aux_Lo          Chr1102x_DualFirBinAux_Lo

Dual Second Binning Aux_Hi         Chr1102x_DualSecBinAux_Hi

Dual Second Binning Aux_Lo         Chr1102x_DualSecBinAux_Lo

First Binning Upper                Chr1102x_FirstBinUpper

First Binning Lower                Chr1102x_FirstBinLower

Trigger Source of Dual             Chr1102x_TriggerSourceofDual

Measure Speed of Dual              Chr1102x_MeasureSpeedofDual

Fetch Result for Dual              Chr1102x_FetchResultforDual

Fetch

Fetch Result                          Chr1102x_FetchResult

Init

Set Initiate Immediate                Chr1102x_SetInitiateImmediate

Set Initiate Continuous               Chr1102x_SetInitiateContinuous

Query

Query Initiate Continuous          Chr1102x_Q_IntContinus

SENSE

Set Average Count                     Chr1102x_SetAverageCount

Set Correction Procedure              Chr1102x_SetCorrectionProcedure

Set Correction Data of Open           Chr1102x_SetCorrectionDataOfOpen

Set Correction Data of Short          Chr1102x_SetCorrectionDataOfShort

Set Correction Data Short 1 Ohm       Chr1102x_SetCorrectionDataOfShort1Ohm

Set Corr_Data Short 100MOhm           Chr1102x_SetCorr_DataShort100MOhm

Set Measurement Speed                 Chr1102x_SetMeasurementSpeed

Set Auto Range State                  Chr1102x_SetAutoRangeState

Set Measure Range                     Chr1102x_SetMeasureRange

Set Measure Transformer State         Chr1102x_SetMeasureTransformerState

Set L2 A_B Parameter                  Chr1102x_SetL2A_BParameter

Set Measure Parallel Mode             Chr1102x_SetMeasureParallelMode

Query

Query Average Count                Chr1102x_Q_Average

Query Correction Value             Chr1102x_QueryCorrectionValue

Query Correction Data of Open      Chr1102x_Q_CorrectionDataofopen

Query Correction Data of Short     Chr1102x_Q_Corrctiondataofshort

Query Corr Data of Short 1 Ohm     Chr1102x_Q_correctiondataofshort1ohm

Query Corr Data Short 100MOhm      Chr1102x_Q_Corrdatashrot100oh

Query Measurement Speed            Chr1102x_Q_measurementspeed

Query Auto Range State             Chr1102x_Q_Autorangestate

Query Measure Range                Chr1102x_Q_Measurerange

Query Measure Transformer State    Chr1102x_Q_Meatranformerstate

Query L2 A_B Parameter             Chr1102x_Q_L2ABParameter

Query Measure Parallel Mode        Chr1102x_Q_Measureparallelmode

SOURCE

Set Test Frequency                    Chr1102x_SetTestFrequency

Set Test Voltage                      Chr1102x_SetTestVoltage

Set BIAS Current State                Chr1102x_SetBIASCurrentState

Set Voltage Offset State              Chr1102x_SetVoltageOffsetstate

Set BIAS Current Value                Chr1102x_SetBIASCurrentValue

Query

Query Test Frequency               Chr1102x_Q_Testfrequency

Query Test Voltage                 Chr1102x_Q_Testvoltage

Query BIAS Current State           Chr1102x_Q_BIASCurrentState

Query Voltage Offset State         Chr1102x_Q_VoltageOffsetState

Query BIAS Current Value           Chr1102x_Q_BIAScurrentvalue

SYSTEM

Set Beeper                            Chr1102x_SetBeeper

Set Beeper State                      Chr1102x_SetBeeperState

Set Key Lock State                    Chr1102x_SetKeyLockState

Set Back to Default State             Chr1102x_SetBacktoDefaultState

Set Ouput Impedance                   Chr1102x_SetOuputImpedance

Set System Integration                Chr1102x_SetSystemIntegration

Set Buzzer Action Type                Chr1102x_SetBuzzerActionType

Set Handler Mode                      Chr1102x_SetHandlerMode

Set LINK 1320 Enable                  Chr1102x_SetLINK1320Enable

Query

Query Beeper State                 Chr1102x_Q_BeeperState

Query Key Lock State               Chr1102x_Q_Keylockstate

Query Version                      Chr1102x_QueryVersion

Query Ouput Impedance              Chr1102x_Q_Oupputimpedance

Query System Integration           Chr1102x_Q_Systemintegration

Query Buzzer Action Type           Chr1102x_Q_Buzzeractiontype

Query Handler Mode                 Chr1102x_Q_Handermode

Query LINK 1320 Enable State       Chr1102x_Q_LINKSTATE

TRIGGER

Set Trigger Delay Time                Chr1102x_SetTriggerDelayTime

Set Trigger Immediate                 Chr1102x_SetTriggerState

Set Trigger Source                    Chr1102x_SetTriggerSource

Set Trigger Edge                      Chr1102x_SetTriggerEdge

Query

Query Trigger Delay Time           Chr1102x_Q_TriggerDelayTime

Query Trigger Source               Chr1102x_Q_TriggerSource

Query Trigger Edge                 Chr1102x_Q_Triggeredge

STATUS

Set Enable Register Status            Chr1102x_SetEnableRegisterStatus

Query

Quey Status Operation              Chr1102x_QueyStatusOperation

Quey Status Operation Condition    Chr1102x_QueyStatusOperationCondition

Query Enable Register Status       Chr1102x_Q_enableregisterstatus

Configuration Functions

Set/Get/Check Attribute

Set Attribute

Set Attribute ViInt32           Chr1102x_SetAttributeViInt32

Set Attribute ViReal64          Chr1102x_SetAttributeViReal64

Set Attribute ViString          Chr1102x_SetAttributeViString

Set Attribute ViBoolean         Chr1102x_SetAttributeViBoolean

Set Attribute ViSession         Chr1102x_SetAttributeViSession

Get Attribute

Get Attribute ViInt32           Chr1102x_GetAttributeViInt32

Get Attribute ViReal64          Chr1102x_GetAttributeViReal64

Get Attribute ViString          Chr1102x_GetAttributeViString

Get Attribute ViBoolean         Chr1102x_GetAttributeViBoolean

Get Attribute ViSession         Chr1102x_GetAttributeViSession

Check Attribute

Check Attribute ViInt32         Chr1102x_CheckAttributeViInt32

Check Attribute ViReal64        Chr1102x_CheckAttributeViReal64

Check Attribute ViString        Chr1102x_CheckAttributeViString

Check Attribute ViBoolean       Chr1102x_CheckAttributeViBoolean

Check Attribute ViSession       Chr1102x_CheckAttributeViSession

Utility Functions

Reset                                 Chr1102x_reset

Self-Test                             Chr1102x_self_test

Revision Query                        Chr1102x_revision_query

Error-Query                           Chr1102x_error_query

Error Message                         Chr1102x_error_message

Error Info

Get Error Info                     Chr1102x_GetErrorInfo

Clear Error Info                   Chr1102x_ClearErrorInfo

Coercion Info

Get Next Coercion Record           Chr1102x_GetNextCoercionRecord

Locking

Lock Session                       Chr1102x_LockSession

Unlock Session                     Chr1102x_UnlockSession

Instrument I/O

Write Instrument Data              Chr1102x_WriteInstrData

Read Instrument Data               Chr1102x_ReadInstrData

Close                                    Chr1102x_close

________________________________________________________________________________

Chroma 1102x LCR Meter

This instrument driver contains programming support for the Chroma 1102x

LCR Meter.  This driver has all the functions that IVI and VXIplug&play

require.

Note:  This driver requires the VISA and IVI libraries.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

Chr1102x_CheckAttributeViBoolean

ViStatus Chr1102x_CheckAttributeViBoolean (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViBoolean attributeValue);

Purpose

This function checks the validity of a value you specify for a ViBoolean

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViBoolean

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViBoolean are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViBoolean

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_CheckAttributeViInt32

ViStatus Chr1102x_CheckAttributeViInt32 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViInt32 attributeValue);

Purpose

This function checks the validity of a value you specify for a ViInt32

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViInt32 type.

If you choose to see all IVI attributes, the data types appear

to the right of the attribute names in the list box.

Attributes with data types other than ViInt32 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViInt32

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_CheckAttributeViReal64

ViStatus Chr1102x_CheckAttributeViReal64 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViReal64 attributeValue);

Purpose

This function checks the validity of a value you specify for a ViReal64

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViReal64

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViReal64 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViReal64

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_CheckAttributeViSession

ViStatus Chr1102x_CheckAttributeViSession (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViSession attributeValue);

Purpose

This function checks the validity of a value you specify for a ViSession

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViSession

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViSession are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViSession

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_CheckAttributeViString

ViStatus Chr1102x_CheckAttributeViString (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViChar _VI_FAR attributeValue[]);

Purpose

This function checks the validity of a value you specify for a ViString

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViString

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViString are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViChar[]

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_ClearErrorInfo

ViStatus Chr1102x_ClearErrorInfo (ViSession instrumentHandle);

Purpose

This function clears the error information for the current execution

thread and the IVI session you specify.  If you pass VI_NULL for the

Instrument Handle parameter, this function clears the error information

only for the current execution thread.

The error information includes a primary error code, secondary error

code, and an error elaboration string.  For a particular session, this

information is the same as the values held in the following attributes:

CHR1102X_ATTR_PRIMARY_ERROR

CHR1102X_ATTR_SECONDARY_ERROR

CHR1102X_ATTR_ERROR_ELABORATION

This function sets the primary and secondary error codes to VI_SUCCESS

(0), and sets the error elaboration string to "".

The IVI Library also maintains this error information separately for each

thread.  This is useful if you do not have a session handle to pass to

Chr1102x_ClearErrorInfo or Chr1102x_GetErrorInfo, which occurs when a

call to Chr1102x_init or Chr1102x_InitWithOptions fails.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

This function returns a non-zero status code only when it is unable

to clear the error information for the session or thread.  The

function can return the following status codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFF000E  Invalid session handle.

________________________________________________________________________________

Chr1102x_close

ViStatus Chr1102x_close (ViSession instrumentHandle);

Purpose

This function performs the following operations:

- Closes the instrument I/O session.

- Destroys the instrument driver session and all of its attributes.

- Deallocates any memory resources the driver uses.

Notes:

(1) You must unlock the session before calling Chr1102x_close.

(2) After calling Chr1102x_close, you cannot use the instrument driver

again until you call Chr1102x_init or Chr1102x_InitWithOptions.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_CLS

ViStatus Chr1102x_CLS (ViSession instrumentHandle);

Purpose

It clears the status Byte Register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualAutoFirstMea

ViStatus Chr1102x_DualAutoFirstMea (ViSession instrumentHandle,

ViPBoolean dualAutoOfFirstMeasure);

Purpose

It queries auto state of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualAutoOfFirstMeasure

Variable Type       ViBoolean (passed by reference)

This parameter returns auto state of first measure of the dual

function.

Description:

1:ON

2:OFF

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualAutoSecMea

ViStatus Chr1102x_DualAutoSecMea (ViSession instrumentHandle,

ViPBoolean dualAutoOfSecondMeasure);

Purpose

It queries auto state of second measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualAutoOfSecondMeasure

Variable Type       ViBoolean (passed by reference)

This parameter returns auto state of second measure of the dual

function.

Description:

1:ON

2:OFF

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualBinningMode

ViStatus Chr1102x_DualBinningMode (ViSession instrumentHandle,

ViChar _VI_FAR dualBinningMode[]);

Purpose

It queries binning mode of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualBinningMode

Variable Type       ViChar[]

This parameter returns It sets binning mode of the dual

function.(ABS/PCNT)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualFirBinAux_Lo

ViStatus Chr1102x_DualFirBinAux_Lo (ViSession instrumentHandle,

ViPReal64 dualFirstBinningAux_Lo);

Purpose

It queries first binning aux-Lo value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualFirstBinningAux_Lo

Variable Type       ViReal64 (passed by reference)

This parameter retruns first binning aux-Lo value of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualFirstBinNom

ViStatus Chr1102x_DualFirstBinNom (ViSession instrumentHandle,

ViPReal64 dualFirstBinningNominal);

Purpose

It queries first binning nominal value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualFirstBinningNominal

Variable Type       ViReal64 (passed by reference)

This parameter retruns first binning nominal value of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualFirstMeaRange

ViStatus Chr1102x_DualFirstMeaRange (ViSession instrumentHandle,

ViPReal64 dualFirstMeasureRange);

Purpose

It queries range value of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualFirstMeasureRange

Variable Type       ViReal64 (passed by reference)

This parameter retruns range value of first measure of the dual

function.

Unit:ohm

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualFirstSouFreq

ViStatus Chr1102x_DualFirstSouFreq (ViSession instrumentHandle,

ViPReal64 dualFirstSourceFreq);

Purpose

It queries source frequnecy range of first test of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualFirstSourceFreq

Variable Type       ViReal64 (passed by reference)

This parameter retruns source frequnecy range of first test of the

dual function.

Unit:Hz

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualFriBinAux

ViStatus Chr1102x_DualFriBinAux (ViSession instrumentHandle,

ViPReal64 dualFirstBinningAux_Hi);

Purpose

It queries first binning aux-Hi value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualFirstBinningAux_Hi

Variable Type       ViReal64 (passed by reference)

This parameter retruns first binning aux-Hi value of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualSecBinAux_Hi

ViStatus Chr1102x_DualSecBinAux_Hi (ViSession instrumentHandle,

ViPReal64 dualSecondBinningAux_Hi);

Purpose

It queries second binning aux-Hi value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualSecondBinningAux_Hi

Variable Type       ViReal64 (passed by reference)

This parameter retruns first binning aux-Lo value of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualSecBinAux_Lo

ViStatus Chr1102x_DualSecBinAux_Lo (ViSession instrumentHandle,

ViPReal64 dualSecondBinningAux_Lo);

Purpose

It queries second binning aux-Hi value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualSecondBinningAux_Lo

Variable Type       ViReal64 (passed by reference)

This parameter retruns second binning aux-Hi value of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualSecMeaRange

ViStatus Chr1102x_DualSecMeaRange (ViSession instrumentHandle,

ViPReal64 dualSecondMeasureRange);

Purpose

It queries range value of second measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualSecondMeasureRange

Variable Type       ViReal64 (passed by reference)

This parameter retruns range value of second measure of the dual

function.

Unit:ohm

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_DualSecSouFreq

ViStatus Chr1102x_DualSecSouFreq (ViSession instrumentHandle,

ViPReal64 setDualSecondSourceFreq);

Purpose

It queries source frequnecy range of second test of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualSecondSourceFreq

Variable Type       ViReal64 (passed by reference)

This parameter retruns source frequnecy range of second test of the

dual function.

Unit:Hz

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_error_message

ViStatus Chr1102x_error_message (ViSession instrumentHandle,

ViStatus errorCode,

ViChar _VI_FAR errorMessage[]);

Purpose

This function converts a status code returned by an instrument driver

function into a user-readable string.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

You can pass VI_NULL for this parameter.  This is useful when one of

the initialize functions fail.

Default Value:  VI_NULL

errorCode

Variable Type       ViStatus

Pass the Status parameter that is returned from any of the instrument

driver functions.

Default Value:  0  (VI_SUCCESS)

errorMessage

Variable Type       ViChar[]

Returns the user-readable message string that corresponds to the

status code you specify.

You must pass a ViChar array with at least 256 bytes.

Return Value

Reports the status of this operation.

This function can return only three possible status codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0085  Unknown status code (warning).

BFFF000A  Invalid parameter (Error Message buffer is VI_NULL).

________________________________________________________________________________

Chr1102x_error_query

ViStatus Chr1102x_error_query (ViSession instrumentHandle,

ViPInt32 errorCode,

ViChar _VI_FAR errorMessage[]);

Purpose

This function reads an error code and a message from the instrument's

error queue.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

errorCode

Variable Type       ViInt32 (passed by reference)

Returns the error code read from the instrument's error queue.

errorMessage

Variable Type       ViChar[]

Returns the error message string read from the instrument's error

message queue.

You must pass a ViChar array with at least 256 bytes.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFC0104  Error Query not supported

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_FetchResult

ViStatus Chr1102x_FetchResult (ViSession instrumentHandle,

ViPInt16 measureStatus,

ViPReal64 firstData1,

ViPReal64 firstData2,

ViPInt16 compare1Status,

ViPInt16 compare2Status,

ViPInt32 BINResult);

Purpose

Gets initiate command obtaining values.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

measureStatus

Variable Type       ViInt16 (passed by reference)

First fetch status.

0:Nominal

1:Over load.

2:connection error.

firstData1

Variable Type       ViReal64 (passed by reference)

This parameter returns Master parameter test value

firstData2

Variable Type       ViReal64 (passed by reference)

This parameter returns Secondary parameter test value

compare1Status

Variable Type       ViInt16 (passed by reference)

This parameter returns Master parameter comparison result

When comparator is opening, there are outputs as the following:

0 :Without secondary parameter

1 :Test value within the range

2 :Test value over high

4 :Test value over low

When comparator is close, output as following:

-1:NONE

compare2Status

Variable Type       ViInt16 (passed by reference)

This parameter returns Secondary parameter comparison result

When comparator is opening, there are outputs as the following:

0 :Without secondary parameter

1 :Test value within the range

2 :Test value over high

4 :Test value over low

When comparator is close, output as following:

-1:NONE

BINResult

Variable Type       ViInt32 (passed by reference)

When BINNING is opening, there are outputs as the following.

<BIN> Sequence result

0   :B parameter fails

1~8 :A, B parameter pass

9 A parameter fails

When BIN>9 is opening, there are outputs as the following.

<BIN>9> Sequence result

0    :B parameter fails

1.99 :A, B parameter pass

100  :A parameter fails

When BIN>9 is close, output as following:

-1:NONE

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_FetchResultforDual

ViStatus Chr1102x_FetchResultforDual (ViSession instrumentHandle,

ViPInt16 firstState,

ViPReal64 firstData1,

ViPReal64 firstData2,

ViPInt16 secondState,

ViPReal64 secondData2,

ViPInt16 firstBIN1,

ViPInt16 firstBIN2,

ViPInt16 secondBIN2);

Purpose

Gets initiate command obtaining values.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

firstState

Variable Type       ViInt16 (passed by reference)

First fetch status.

0:Nominal

1:Over load.

2:connection error.

firstData1

Variable Type       ViReal64 (passed by reference)

First fetch test result of primary parameter.

firstData2

Variable Type       ViReal64 (passed by reference)

First fetch test result of secondary parameter.

secondState

Variable Type       ViInt16 (passed by reference)

Second fetch status.

0:Nominal

1:Over load.

2:connection error.

secondData2

Variable Type       ViReal64 (passed by reference)

Second fetch test result of primary parameter.

firstBIN1

Variable Type       ViInt16 (passed by reference)

First fetch BIN result of primary parameter.

0:BIN OUT

1:BIN Number

firstBIN2

Variable Type       ViInt16 (passed by reference)

First fetch BIN result of secondary parameter.

0:NOGO

1:GO

secondBIN2

Variable Type       ViInt16 (passed by reference)

Second fetch BIN result of secondary parameter.

0:NOGO

1:GO

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_FirstBinLower

ViStatus Chr1102x_FirstBinLower (ViSession instrumentHandle,

ViPReal64 firstBinningLower,

ViInt32 binNumber);

Purpose

It queries the lower limit value of each BINs.  When BINning: MODE ABS,

it is the absolute value of inputted upper limit.  When BINning: MODE

PCNT, it is the percentage of inputted upper limit.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

firstBinningLower

Variable Type       ViReal64 (passed by reference)

This parameter retruns the lower limit value of each BINs.  When

BINning: MODE ABS, it is the absolute value of inputted upper limit.

When BINning: MODE PCNT, it is the percentage of inputted upper

limit.

binNumber

Variable Type       ViInt32

It sets the number of BINs of dual function.

Valid Range: 1 ~ 4.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_FirstBinUpper

ViStatus Chr1102x_FirstBinUpper (ViSession instrumentHandle,

ViPReal64 firstBinningUpper,

ViInt32 binNumber);

Purpose

It queries the upper limit value of each BINs.  When BINning: MODE ABS,

it is the absolute value of inputted upper limit.  When BINning: MODE

PCNT, it is the percentage of inputted upper limit.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

firstBinningUpper

Variable Type       ViReal64 (passed by reference)

This parameter retruns the upper limit value of each BINs.  When

BINning: MODE ABS, it is the absolute value of inputted upper limit.

When BINning: MODE PCNT, it is the percentage of inputted upper

limit.

binNumber

Variable Type       ViInt32

It sets the number of BINs of dual function.

Valid Range: 1 ~ 4.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_GetAttributeViBoolean

ViStatus Chr1102x_GetAttributeViBoolean (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPBoolean attributeValue);

Purpose

This function queries the value of a ViBoolean attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViBoolean

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViBoolean are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViBoolean (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViBoolean variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_GetAttributeViInt32

ViStatus Chr1102x_GetAttributeViInt32 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPInt32 attributeValue);

Purpose

This function queries the value of a ViInt32 attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViInt32 type.

If you choose to see all IVI attributes, the data types appear

to the right of the attribute names in the list box.

Attributes with data types other than ViInt32 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViInt32 (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViInt32 variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_GetAttributeViReal64

ViStatus Chr1102x_GetAttributeViReal64 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPReal64 attributeValue);

Purpose

This function queries the value of a ViReal64 attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViReal64

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViReal64 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViReal64 (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViReal64 variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_GetAttributeViSession

ViStatus Chr1102x_GetAttributeViSession (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPSession attributeValue);

Purpose

This function queries the value of a ViSession attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViSession

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViSession are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViSession (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViSession variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_GetAttributeViString

ViStatus Chr1102x_GetAttributeViString (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViInt32 bufferSize,

ViChar _VI_FAR attributeValue[]);

Purpose

This function queries the value of a ViString attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

You must provide a ViChar array to serve as a buffer for the value.  You

pass the number of bytes in the buffer as the Buffer Size parameter.  If

the current value of the attribute, including the terminating NUL byte,

is larger than the size you indicate in the Buffer Size parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an ASCII

NUL byte at the end of the buffer, and returns the buffer size you must

pass to get the entire value.  For example, if the value is "123456" and

the Buffer Size is 4, the function places "123" into the buffer and

returns 7.

If you want to call this function just to get the required buffer size,

you can pass 0 for the Buffer Size and VI_NULL for the Attribute Value

buffer.

If you want the function to fill in the buffer regardless of the   number

of bytes in the value, pass a negative number for the Buffer Size

parameter.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViString

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViString are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Attribute Value parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Attribute Value buffer

parameter.

attributeValue

Variable Type       ViChar[]

The buffer in which the function returns the current value of the

attribute.  The buffer must be of type ViChar and have at least as

many bytes as indicated in the Buffer Size parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you specify 0 for the Buffer Size parameter, you can pass VI_NULL

for this parameter.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Reports the status of this operation.

If the function succeeds and the buffer you pass is large enough to

hold the entire value, the function returns 0.

If the current value of the attribute, including the terminating NUL

byte, is larger than the size you indicate in the Buffer Size

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value.  For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

If the function fails for some other reason, it returns a negative

error code.  For more information on error codes, refer to the Status

return value control in one of the other function panels.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_GetBinningResult

ViStatus Chr1102x_GetBinningResult (ViSession instrumentHandle,

ViPInt32 binningResult);

Purpose

It queries the result of BINNING, if BINNING is OFF or sequence result is

0, output +0.  If the sequence result is OUT, output +9.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binningResult

Variable Type       ViInt32 (passed by reference)

It queries the result of BINNING, if BINNING is OFF or sequence

result is 0, output +0.  If the sequence result is OUT, output +9.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_GetErrorInfo

ViStatus Chr1102x_GetErrorInfo (ViSession instrumentHandle,

ViPStatus primaryError,

ViPStatus secondaryError,

ViChar _VI_FAR errorElaboration[]);

Purpose

This function returns the error information associated with an IVI

session or with the current execution thread.  If you specify a valid IVI

session for the Instrument Handle parameter, this function retrieves and

then clears the error information for the session.  If you pass VI_NULL

for the Instrument Handle parameter, this function retrieves and then

clears the error information for the current execution thread.

The error information includes a primary error, secondary error, and an

error elaboration string.  For a particular session, this information is

the same as the values held in the following attributes:

CHR1102X_ATTR_PRIMARY_ERROR

CHR1102X_ATTR_SECONDARY_ERROR

CHR1102X_ATTR_ERROR_ELABORATION

The IVI Library also maintains this error information separately for each

thread.  This is useful if you do not have a session handle to pass to

Chr1102x_GetErrorInfo or Chr1102x_ClearErrorInfo, which

occurs when a call to Chr1102x_init or Chr1102x_InitWithOptions fails.

You can call Chr1102x_error_message to obtain a text description of the

primary or secondary error value.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  VI_NULL

primaryError

Variable Type       ViStatus (passed by reference)

The primary error code for the session or execution thread.   The

value is the same as the one held by the attribute

CHR1102X_ATTR_PRIMARY_ERROR.

A value of VI_SUCCESS (0) indicates that no error occurred,  A

positive value indicates a warning.  A negative value indicates an

error.

You can call Chr1102x_error_message to get a text description of the

value.

If you are not interested in this value, you can pass VI_NULL.

secondaryError

Variable Type       ViStatus (passed by reference)

The secondary error code for the session or execution thread.  If the

primary error code is non-zero, the secondary error code can further

describe the error or warning condition.  The value is the same as

the one held by the attribute CHR1102X_ATTR_SECONDARY_ERROR.

A value of VI_SUCCESS (0) indicates no further description.

You can call Chr1102x_error_message to get a text description of the

value.

If you are not interested in this value, you can pass VI_NULL.

errorElaboration

Variable Type       ViChar[]

The error elaboration string for the session or execution thread.  If

the primary error code is non-zero, the elaboration string can

further describe the error or warning condition.   The value is the

same as the one held by the attribute

CHR1102X_ATTR_ERROR_ELABORATION.

If you are not interested in this value, you can pass VI_NULL.

Otherwise, you must pass a ViChar array with at least 256 bytes.

Return Value

Reports the status of this operation.

This function returns a non-zero status code only when it is unable

to obtain the error information for the session or thread.  The

function can return the following status codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0021  Unable to allocate system resource.

BFFF000E  Invalid session handle.

________________________________________________________________________________

Chr1102x_GetNextCoercionRecord

ViStatus Chr1102x_GetNextCoercionRecord (ViSession instrumentHandle,

ViInt32 bufferSize,

ViChar _VI_FAR coercionRecord[]);

Purpose

This function returns the coercion information associated with the IVI

session.  This function retrieves and clears the oldest instance in which

the instrument driver coerced a value you specified to another value.

If you set the CHR1102X_ATTR_RECORD_COERCIONS attribute to VI_TRUE, the

instrument driver keeps a list of all coercions it makes on ViInt32 or

ViReal64 values you pass to instrument driver functions.  You use this

function to retrieve information from that list.

If the next coercion record string, including the terminating NUL byte,

contains more bytes than you indicate in this parameter, the function

copies Buffer Size - 1 bytes into the buffer, places an ASCII NUL byte at

the end of the buffer, and returns the buffer size you must pass to get

the entire value.  For example, if the value is "123456" and the Buffer

Size is 4, the function places "123" into the buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Coercion Record buffer

parameter.

The function returns an empty string in the Coercion Record parameter if

no coercion records remain for the session.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init function.

The handle identifies a particular instrument session.

Default Value:  None

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Coercion Record parameter.

If the next coercion record string, including the terminating NUL

byte, contains more bytes than you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Coercion Record buffer

parameter.

Default Value:  None

coercionRecord

Variable Type       ViChar[]

Returns the next coercion record for the IVI session.  If there are

no coercion records, the function returns an empty string.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.  If the next coercion record

string, including the terminating NUL byte, contains more bytes than

you indicate with the Buffer Size parameter, the function copies

Buffer Size - 1 bytes into the buffer, places an ASCII NUL byte at

the end of the buffer, and returns the buffer size you must pass to

get the entire value.  For example, if the value is "123456" and the

Buffer Size is 4, the function places "123" into the buffer and

returns 7.

This parameter returns an empty string if no coercion records remain

for the session.

Return Value

Reports the status of this operation.

If the function succeeds and the buffer you pass is large enough to

hold the entire coercion record string, the function returns 0.

If the next coercion record string, including the terminating NUL

byte, is larger than the size you indicate in the Buffer Size

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value.  For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

If the function fails for some other reason, it returns a negative

error code.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_init

ViStatus Chr1102x_init (ViRsrc resourceName, ViBoolean IDQuery,

ViBoolean resetDevice,

ViPSession instrumentHandle, ViInt32 baudRate);

Purpose

This function performs the following initialization actions:

- Creates a new IVI instrument driver session.

- Opens a session to the specified device using the interface and address

you specify for the Resource Name parameter.

- If the ID Query parameter is set to VI_TRUE, this function queries the

instrument ID and checks that it is valid for this instrument driver.

- If the Reset parameter is set to VI_TRUE, this function resets the

instrument to a known state.

- Sends initialization commands to set the instrument to the state

necessary for the operation of the instrument driver.

- Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Note:  This function creates a new session each time you invoke it.

Although you can open more than one IVI session for the same resource, it

is best not to do so.  You can use the same session in multiple program

threads.  You can use the Chr1102x_LockSession and Chr1102x_UnlockSession

functions to protect sections of code that require exclusive access to

the resource.

Parameter List

resourceName

Variable Type       ViRsrc

Pass the resource name of the device to initialize.

You can also pass the name of a virtual instrument or logical name

that you configure with the IVI Configuration utility.  The virtual

instrument identifies a specific device and specifies the initial

settings for the session.  A logical Name identifies a particular

virtual instrument.

Refer to the following table below for the exact grammar to use for

this parameter.  Optional fields are shown in square brackets ([]).

Syntax

------------------------------------------------------

GPIB[board]::<primary address>[::secondary address]::INSTR

VXI[board]::<logical address>::INSTR

GPIB-VXI[board]::<logical address>::INSTR

ASRL<port>::INSTR

<LogicalName>

[VInstr->]<VirtualInstrument>

If you do not specify a value for an optional field, the following

values are used:

Optional Field - Value

------------------------------------------------------

board - 0

secondary address - none (31)

The following table contains example valid values for this parameter.

"Valid Value" - Description

------------------------------------------------------

"GPIB::22::INSTR" - GPIB board 0, primary address 22 no

secondary address

"GPIB::22::5::INSTR" - GPIB board 0, primary address 22

secondary address 5

"GPIB1::22::5::INSTR" - GPIB board 1, primary address 22

secondary address 5

"VXI::64::INSTR" - VXI board 0, logical address 64

"VXI1::64::INSTR" - VXI board 1, logical address 64

"GPIB-VXI::64::INSTR" - GPIB-VXI board 0, logical address 64

"GPIB-VXI1::64::INSTR" - GPIB-VXI board 1, logical address 64

"ASRL2::INSTR" - COM port 2

"SampleInstr" - Logical name "SampleInstr"

"VInstr->xyz432" - Virtual Instrument "xyz432"

"xyz432" - Logical Name or Virtual Instrument "xyz432"

/*=CHANGE:===================================================*

Modify the following default value so that it reflects the default

address for your instrument.  You must make the corresponding change

to the Default Value entry for the control.

*================================================END=CHANGE=*/

Default Value:  "GPIB::17::INSTR"

Note: You specify the resource name with the "VInstr->" if you have

the logical name that is the same as the virtual instrument name and

you want to explicitly use the virtual instrument name. Otherwise,

the driver uses the logical name.

IDQuery

Variable Type       ViBoolean

Specify whether you want the instrument driver to perform an ID

Query.

Valid Range:

VI_TRUE  (1) - Perform ID Query (Default Value)

VI_FALSE (0) - Skip ID Query

When you set this parameter to VI_TRUE, the driver verifies that the

instrument you initialize is a type that this driver supports.

Circumstances can arise where it is undesirable to send an ID Query

command string to the instrument.  When you set this parameter to

VI_FALSE, the function initializes the instrument without performing

an ID Query.

resetDevice

Variable Type       ViBoolean

Specify whether you want the to reset the instrument during the

initialization procedure.

Valid Range:

VI_TRUE  (1) - Reset Device (Default Value)

VI_FALSE (0) - Don't Reset

instrumentHandle

Variable Type       ViSession (passed by reference)

Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Notes:

(1) This function creates a new session each time you invoke it.

This is useful if you have multiple physical instances of the same

type of instrument.

(2) Avoid creating multiple concurrent sessions to the same physical

instrument.  Although you can create more than one IVI session for

the same resource, it is best not to do so.  A better approach is to

use the same IVI session in multiple execution threads.  You can use

functions Chr1102x_LockSession and Chr1102x_UnlockSession to protect

sections of code that require exclusive access to the resource.

baudRate

Variable Type       ViInt32

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFC0101  ID Query not supported.

3FFC0102  Reset not supported.

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFA0000  Cannot load IVI engine.

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFA0054  Bad channel name in Channel List.

BFFC0011  Instrument returned invalid response to ID Query.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_InitWithOptions

ViStatus Chr1102x_InitWithOptions (ViRsrc resourceName,

ViBoolean IDQuery,

ViBoolean resetDevice,

ViString optionString,

ViPSession instrumentHandle);

Purpose

This function performs the following initialization actions:

- Creates a new IVI instrument driver and optionally sets the initial

state of the following session attributes:

CHR1102X_ATTR_RANGE_CHECK

CHR1102X_ATTR_QUERY_INSTR_STATUS

CHR1102X_ATTR_CACHE

CHR1102X_ATTR_SIMULATE

CHR1102X_ATTR_RECORD_COERCIONS

- Opens a session to the specified device using the interface and address

you specify for the Resource Name parameter.

- If the ID Query parameter is set to VI_TRUE, this function queries the

instrument ID and checks that it is valid for this instrument driver.

- If the Reset parameter is set to VI_TRUE, this function resets the

instrument to a known state.

- Sends initialization commands to set the instrument to the state

necessary for the operation of the instrument driver.

- Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Note:  This function creates a new session each time you invoke it.

Although you can open more than one IVI session for the same resource, it

is best not to do so.  You can use the same session in multiple program

threads.  You can use the Chr1102x_LockSession and Chr1102x_UnlockSession

functions to protect sections of code that require exclusive access to

the resource.

Parameter List

resourceName

Variable Type       ViRsrc

Pass the resource name of the device to initialize.

You can also pass the name of a virtual instrument or logical name

that you configure with the IVI Configuration utility.  The virtual

instrument identifies a specific device and specifies the initial

settings for the session.  A logical Name identifies a particular

virtual instrument.

Refer to the following table below for the exact grammar to use for

this parameter.  Optional fields are shown in square brackets ([]).

Syntax

------------------------------------------------------

GPIB[board]::<primary address>[::secondary address]::INSTR

VXI[board]::<logical address>::INSTR

GPIB-VXI[board]::<logical address>::INSTR

ASRL<port>::INSTR

<LogicalName>

[VInstr->]<VirtualInstrument>

If you do not specify a value for an optional field, the following

values are used:

Optional Field - Value

------------------------------------------------------

board - 0

secondary address - none (31)

The following table contains example valid values for this parameter.

"Valid Value" - Description

------------------------------------------------------

"GPIB::22::INSTR" - GPIB board 0, primary address 22 no

secondary address

"GPIB::22::5::INSTR" - GPIB board 0, primary address 22

secondary address 5

"GPIB1::22::5::INSTR" - GPIB board 1, primary address 22

secondary address 5

"VXI::64::INSTR" - VXI board 0, logical address 64

"VXI1::64::INSTR" - VXI board 1, logical address 64

"GPIB-VXI::64::INSTR" - GPIB-VXI board 0, logical address 64

"GPIB-VXI1::64::INSTR" - GPIB-VXI board 1, logical address 64

"ASRL2::INSTR" - COM port 2

"SampleInstr" - Logical name "SampleInstr"

"VInstr->xyz432" - Virtual Instrument "xyz432"

"xyz432" - Logical Name or Virtual Instrument "xyz432"

/*=CHANGE:===================================================*

Modify the following default value so that it reflects the default

address for your instrument.  You must make the corresponding change

to the Default Value entry for the control.

*================================================END=CHANGE=*/

Default Value:  "GPIB::17::INSTR"

Note: You specify the resource name with the "VInstr->" if you have

the logical name that is the same as the virtual instrument name and

you want to explicitly use the virtual instrument name. Otherwise,

the driver uses the logical name.

IDQuery

Variable Type       ViBoolean

Specify whether you want the instrument driver to perform an ID

Query.

Valid Range:

VI_TRUE  (1) - Perform ID Query (Default Value)

VI_FALSE (0) - Skip ID Query

When you set this parameter to VI_TRUE, the driver verifies that the

instrument you initialize is a type that this driver supports.

Circumstances can arise where it is undesirable to send an ID Query

command string to the instrument.  When you set this parameter to

VI_FALSE, the function initializes the instrument without performing

an ID Query.

resetDevice

Variable Type       ViBoolean

Specify whether you want the to reset the instrument during the

initialization procedure.

Valid Range:

VI_TRUE  (1) - Reset Device (Default Value)

VI_FALSE (0) - Don't Reset

optionString

Variable Type       ViString

You can use this control to set the initial value of certain

attributes for the session.  The following table lists the attributes

and the name you use in this parameter to identify the attribute.

Name              Attribute Defined Constant

--------------------------------------------

RangeCheck        CHR1102X_ATTR_RANGE_CHECK

QueryInstrStatus  CHR1102X_ATTR_QUERY_INSTR_STATUS

Cache             CHR1102X_ATTR_CACHE

Simulate          CHR1102X_ATTR_SIMULATE

RecordCoercions   CHR1102X_ATTR_RECORD_COERCIONS

The format of this string is, "AttributeName=Value" where

AttributeName is the name of the attribute and Value is the value to

which the attribute will be set.  To set multiple attributes,

separate their assignments with a comma.

If you pass NULL or an empty string for this parameter, the session

uses the default values for the attributes.   You can override the

default values by assigning a value explicitly in a string you pass

for this parameter.  You do not have to specify all of the attributes

and may leave any of them out.  If you do not specify one of the

attributes, its default value will be used.

The default values for the attributes are shown below:

Attribute Name     Default Value

----------------   -------------

RangeCheck         VI_TRUE

QueryInstrStatus   VI_TRUE

Cache              VI_TRUE

Simulate           VI_FALSE

RecordCoercions    VI_FALSE

The following are the valid values for ViBoolean attributes:

True:     1, True, or VI_TRUE

False:    0, False, or VI_FALSE

Default Value:

"Simulate=0,RangeCheck=1,QueryInstrStatus=1,Cache=1"

instrumentHandle

Variable Type       ViSession (passed by reference)

Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Notes:

(1) This function creates a new session each time you invoke it.

This is useful if you have multiple physical instances of the same

type of instrument.

(2) Avoid creating multiple concurrent sessions to the same physical

instrument.  Although you can create more than one IVI session for

the same resource, it is best not to do so.  A better approach is to

use the same IVI session in multiple execution threads.  You can use

functions Chr1102x_LockSession and Chr1102x_UnlockSession to protect

sections of code that require exclusive access to the resource.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFC0101  ID Query not supported.

3FFC0102  Reset not supported.

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFA0000  Cannot load IVI engine.

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFA0049  Missing option name (nothing before the '=').

BFFA004A  Missing option value (nothing after the '=').

BFFA004B  Bad option name.

BFFA004C  Bad option value.

BFFA0054  Bad channel name in Channel List

BFFC0011  Instrument returned invalid response to ID Query.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_LockSession

ViStatus Chr1102x_LockSession (ViSession instrumentHandle,

ViPBoolean callerHasLock);

Purpose

This function obtains a multithread lock on the instrument session.

Before it does so, it waits until all other execution threads have

released their locks on the instrument session.

Other threads might have obtained a lock on this session in the following

ways:

- The user's application called Chr1102x_LockSession.

- A call to the instrument driver locked the session.

- A call to the IVI engine locked the session.

After your call to Chr1102x_LockSession returns successfully, no other

threads can access the instrument session until you call

Chr1102x_UnlockSession.

Use Chr1102x_LockSession and Chr1102x_UnlockSession around a sequence of

calls to instrument driver functions if you require that the instrument

retain its settings through the end of the sequence.

You can safely make nested calls to Chr1102x_LockSession within the same

thread.  To completely unlock the session, you must balance each call to

Chr1102x_LockSession with a call to Chr1102x_UnlockSession.  If, however,

you use the Caller Has Lock parameter in all calls to

Chr1102x_LockSession and Chr1102x_UnlockSession within a function, the

IVI Library locks the session only once within the function regardless of

the number of calls you make to Chr1102x_LockSession.  This allows you to

call Chr1102x_UnlockSession just once at the end of the function.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

callerHasLock

Variable Type       ViBoolean (passed by reference)

This parameter serves as a convenience.  If you do not want to use

this parameter, pass VI_NULL.

Use this parameter in complex functions to keep track of whether you

obtain a lock and therefore need to unlock the session.  Pass the

address of a local ViBoolean variable.  In the declaration of the

local variable, initialize it to VI_FALSE.  Pass the address of the

same local variable to any other calls you make to

Chr1102x_LockSession or Chr1102x_UnlockSession in the same function.

The parameter is an input/output parameter.  Chr1102x_LockSession and

Chr1102x_UnlockSession each inspect the current value and take the

following actions:

- If the value is VI_TRUE, Chr1102x_LockSession does not lock the

session again.  If the value is VI_FALSE, Chr1102x_LockSession

obtains the lock and sets the value of the parameter to VI_TRUE.

- If the value is VI_FALSE, Chr1102x_UnlockSession does not attempt

to unlock the session.  If the value is VI_TRUE,

Chr1102x_UnlockSession releases the lock and sets the value of the

parameter to VI_FALSE.

Thus, you can, call Chr1102x_UnlockSession at the end of your

function without worrying about whether you actually have the lock.

Example:

ViStatus TestFunc (ViSession vi, ViInt32 flags)

{

ViStatus error = VI_SUCCESS;

ViBoolean haveLock = VI_FALSE;

if (flags & BIT_1)

{

viCheckErr( Chr1102x_LockSession(vi, &haveLock));

viCheckErr( TakeAction1(vi));

if (flags & BIT_2)

{

viCheckErr( Chr1102x_UnlockSession(vi, &haveLock));

viCheckErr( TakeAction2(vi));

viCheckErr( Chr1102x_LockSession(vi, &haveLock);

}

if (flags & BIT_3)

viCheckErr( TakeAction3(vi));

}

Error:

/*

At this point, you cannot really be sure that

you have the lock.  Fortunately, the haveLock

variable takes care of that for you.

*/

Chr1102x_UnlockSession(vi, &haveLock);

return error;

}

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_MeasureSpeedofDual

ViStatus Chr1102x_MeasureSpeedofDual (ViSession instrumentHandle,

ViPReal64 measureSpeedOfDual);

Purpose

It queries measure speed of dual mode.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

measureSpeedOfDual

Variable Type       ViReal64 (passed by reference)

This parameter retruns measure speed of dual mode

Fast:0.025

Medium:0.065

Slow:0.5

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Autorangestate

ViStatus Chr1102x_Q_Autorangestate (ViSession instrumentHandle,

ViPBoolean autoRangeState);

Purpose

It queries whether turn on auto file selection mode.

1:On

0:Off

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

autoRangeState

Variable Type       ViBoolean (passed by reference)

This parameter returns whether turn on auto file selection mode.

1:On

0:Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Average

ViStatus Chr1102x_Q_Average (ViSession instrumentHandle,

ViChar _VI_FAR averageTimes[]);

Purpose

It queries average times.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

averageTimes

Variable Type       ViChar[]

This parameter returns average times.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_BeeperState

ViStatus Chr1102x_Q_BeeperState (ViSession instrumentHandle,

ViPInt32 beeperState);

Purpose

It sets whether buzzer is turned on or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

beeperState

Variable Type       ViInt32 (passed by reference)

This parameter returns whether buzzer is turned on or not.

0:OFF - turn off buzzer

1:ON - turn on buzzer (small)

2:LARGE - turn on buzzer (large)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_BIASCurrentState

ViStatus Chr1102x_Q_BIASCurrentState (ViSession instrumentHandle,

ViPBoolean BIASCurrentState);

Purpose

It queries the output condition of bias current source.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BIASCurrentState

Variable Type       ViBoolean (passed by reference)

This parameter returns the output condition of bias current source.

1:On

0:Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_BIAScurrentvalue

ViStatus Chr1102x_Q_BIAScurrentvalue (ViSession instrumentHandle,

ViPReal64 BIASCurrentValue);

Purpose

It returns and sets bias current value.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BIASCurrentValue

Variable Type       ViReal64 (passed by reference)

This parameter returns bias current value.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_BINHI9

ViStatus Chr1102x_Q_BINHI9 (ViSession instrumentHandle,

ViPReal64 BINHILimit);

Purpose

It queries the upper limit value of BIN NO when BIN NO>9 function in

SYSTEM CONFIG is enabled.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BINHILimit

Variable Type       ViReal64 (passed by reference)

This parameter retruns the upper limit value of BIN NO when BIN NO>9

function in SYSTEM CONFIG is enabled.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_BINLo9

ViStatus Chr1102x_Q_BINLo9 (ViSession instrumentHandle,

ViPReal64 BINLoLimit);

Purpose

It queries the lower limit value of BIN NO when BIN NO>9 function in

SYSTEM CONFIG is enabled.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BINLoLimit

Variable Type       ViReal64 (passed by reference)

This parameter retruns the lower limit value of BIN NO when BIN NO>9

function in SYSTEM CONFIG is enabled.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_BinningState

ViStatus Chr1102x_Q_BinningState (ViSession instrumentHandle,

ViPBoolean binningState);

Purpose

It queries BINNING is open or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binningState

Variable Type       ViBoolean (passed by reference)

This parameter returns BINNING is open or not.

1: ON (open)

0: OFF (close)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Buzzer_Outdef

ViStatus Chr1102x_Q_Buzzer_Outdef (ViSession instrumentHandle,

ViChar _VI_FAR buzzerOutputDefine[],

ViBoolean calculateNumber);

Purpose

It queries Buzzer output definition.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

buzzerOutputDefine

Variable Type       ViChar[]

This parameter returns Buzzer output definition.

Description:

FAIL(0):When the comparison result is FAIL, buzzer will be action.

PASS(1):When the comparison result is PASS, buzzer will be action.

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Buzzeractiontype

ViStatus Chr1102x_Q_Buzzeractiontype (ViSession instrumentHandle,

ViChar _VI_FAR buzzerActionType[]);

Purpose

It queries buzzer action is PULSE or CONTINUOUS when BINSORTING and

COMPARE.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

buzzerActionType

Variable Type       ViChar[]

This parameter returns buzzer action is PULSE or CONTINUOUS when

BINSORTING and COMPARE.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_BuzzerOutAct

ViStatus Chr1102x_Q_BuzzerOutAct (ViSession instrumentHandle,

ViPBoolean buzzerOutputEnable,

ViBoolean calculateNumber);

Purpose

It queries whether buzzer is action or not.

Description:

ON (1):    Turn on buzzer.

OFF (0):   Turn off buzzer.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

buzzerOutputEnable

Variable Type       ViBoolean (passed by reference)

This parameter returns whether buzzer is action or not.

Description:

ON (1):    Turn on buzzer.

OFF (0):   Turn off buzzer.

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Calc1DualFirst

ViStatus Chr1102x_Q_Calc1DualFirst (ViSession instrumentHandle,

ViChar _VI_FAR calc1OfDualFirstMeasure[]);

Purpose

It queries master parameter of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calc1OfDualFirstMeasure

Variable Type       ViChar[]

This parameter returns master parameter of first measure of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Calc2Dual2nd

ViStatus Chr1102x_Q_Calc2Dual2nd (ViSession instrumentHandle,

ViChar _VI_FAR calc2OfDual2ndMeasure[]);

Purpose

It queries secondary parameter of secondary measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calc2OfDual2ndMeasure

Variable Type       ViChar[]

This parameter returns secondary parameter of secondary measure of

the dual function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Calc2DualFirst

ViStatus Chr1102x_Q_Calc2DualFirst (ViSession instrumentHandle,

ViChar _VI_FAR calc2OfDualFirstMeasure[]);

Purpose

It queries secondary parameter of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calc2OfDualFirstMeasure

Variable Type       ViChar[]

This parameter returns master parameter of first measure of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_calcpath

ViStatus Chr1102x_Q_calcpath (ViSession instrumentHandle,

ViChar _VI_FAR calcSubsystem[],

ViBoolean calculateNumber);

Purpose

It queries According to the sequence of performing CALCulate subsystem

return CALCulate subsystem.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calcSubsystem

Variable Type       ViChar[]

This parameter returns According to the sequence of performing

CALCulate subsystem return CALCulate subsystem.

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_compareLimitAction

ViStatus Chr1102x_Q_compareLimitAction (ViSession instrumentHandle,

ViPBoolean compareLimitAction,

ViBoolean calculateNumber);

Purpose

It queries whether the comparison function is action or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

compareLimitAction

Variable Type       ViBoolean (passed by reference)

This parameter returns whether the comparison function is action or

not.

Description:

ON(1)     Turn on comparison function.

OFF(0)    Turn off comparison function.

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Corrctiondataofshort

ViStatus Chr1102x_Q_Corrctiondataofshort (ViSession instrumentHandle,

ViChar _VI_FAR correctionDataOfShort[]);

Purpose

It queries correction value of SHORT >=10[,, that is R, X

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionDataOfShort

Variable Type       ViChar[]

This parameter returns correction value of SHORT >=10[,that is R, X

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Corrdatashrot100oh

ViStatus Chr1102x_Q_Corrdatashrot100oh (ViSession instrumentHandle,

ViChar _VI_FAR corrdatashort100ohm[]);

Purpose

It queries SHORT correction value of RANGE 100m[, that is R, X

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

corrdatashort100ohm

Variable Type       ViChar[]

This parameter SHORT correction value of RANGE 100m[, that is R, X

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_CorrectionDataofopen

ViStatus Chr1102x_Q_CorrectionDataofopen (ViSession instrumentHandle,

ViChar _VI_FAR correctionDataOfOpen[]);

Purpose

It queries OPEN correction value, that is G, B.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionDataOfOpen

Variable Type       ViChar[]

This parameter returns OPEN correction value, that is G, B.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_correctiondataofshort1ohm

ViStatus Chr1102x_Q_correctiondataofshort1ohm

(ViSession instrumentHandle,

ViChar _VI_FAR correctionDataOfShort1Ohm[]);

Purpose

It queries SHORT correction value of RANGE 1[, that is R, X

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionDataOfShort1Ohm

Variable Type       ViChar[]

This parameter returns SHORT correction value of RANGE 1[, that is

R, X

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_enableregisterstatus

ViStatus Chr1102x_Q_enableregisterstatus (ViSession instrumentHandle,

ViPInt32 enableRegisterStatus);

Purpose

It queries enable register status.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

enableRegisterStatus

Variable Type       ViInt32 (passed by reference)

This paremeter returns enable register status.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_ESE

ViStatus Chr1102x_Q_ESE (ViSession instrumentHandle, ViPInt32 ESEValue);

Purpose

It queries start bit in register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

ESEValue

Variable Type       ViInt32 (passed by reference)

This parameter returns start bit in register.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_ESR

ViStatus Chr1102x_Q_ESR (ViSession instrumentHandle, ViPInt32 ESRValue);

Purpose

It queries the contents of the Standard Event Status Register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

ESRValue

Variable Type       ViInt32 (passed by reference)

It queries the contents of the Standard Event Status Register.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Handermode

ViStatus Chr1102x_Q_Handermode (ViSession instrumentHandle,

ViChar _VI_FAR handlerMode[]);

Purpose

It queries HANDLER MODE.The factory default value is CLEAR.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

handlerMode

Variable Type       ViChar[]

This parameter returns HANDLER MODE.The factory default value is

CLEAR.

CLEAR : When using Handler interface, each time before measurement

the last test result will be cleared (Pass or Fail).

HOLD : When using Handler interface, test result (Pass or Fail) will

remain until next test

result is displayed.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_IntContinus

ViStatus Chr1102x_Q_IntContinus (ViSession instrumentHandle,

ViPBoolean initiateContinuousState);

Purpose

It queries whether initial trigger system sequentially.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

initiateContinuousState

Variable Type       ViBoolean (passed by reference)

This parameter returns whether initial trigger system sequentially.

0(Off) :No sequential initial trigger system

1(On)  :With sequential initial trigger system

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Keylockstate

ViStatus Chr1102x_Q_Keylockstate (ViSession instrumentHandle,

ViPBoolean keyLockState);

Purpose

It queries whether key lock or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

keyLockState

Variable Type       ViBoolean (passed by reference)

This paremeter returns whether key lock or not.

0:Off (Unlock)

1:On  (Key Lock)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_L2ABParameter

ViStatus Chr1102x_Q_L2ABParameter (ViSession instrumentHandle,

ViPBoolean queryL2A_BParameter);

Purpose

It queries the parameters of L2A and L2B.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

queryL2A_BParameter

Variable Type       ViBoolean (passed by reference)

This parameter returns the parameters of L2A and L2B.

1:L2A

0:L2B

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_LCDDisplay

ViStatus Chr1102x_Q_LCDDisplay (ViSession instrumentHandle,

ViPBoolean LCDDisplayStatus);

Purpose

It queries LCD display is turned on or turn off status.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LCDDisplayStatus

Variable Type       ViBoolean (passed by reference)

This parameter returns LCD Display Status

Description:

ON (1)

OFF (0)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_LCDPage

ViStatus Chr1102x_Q_LCDPage (ViSession instrumentHandle,

ViPInt32 LCDText1Page);

Purpose

It queries display screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LCDText1Page

Variable Type       ViInt32 (passed by reference)

This parameter returns display screen.

Description:

1: Measuring screen

2: Comparison screen

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_LCDText2Page

ViStatus Chr1102x_Q_LCDText2Page (ViSession instrumentHandle,

ViPInt32 LCDText2Page);

Purpose

It queries page of measuring screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LCDText2Page

Variable Type       ViInt32 (passed by reference)

This parameter returns page of measuring screen.

Description:

1: Show the first page of measuring screen (test frequency, level and

parameter).

2: Show the second page of measuring screen (file, average time and

trigger mode).

3: Show the third page of measuring screen (DC bias, binning and

comparison setting).

4:Show the four page of measuring screen (bias I state,Bias parameter

setting).

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_lengthoftest

ViStatus Chr1102x_Q_lengthoftest (ViSession instrumentHandle,

ViPInt32 lengthOfTestCable);

Purpose

It queries length of test cable.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

lengthOfTestCable

Variable Type       ViInt32 (passed by reference)

This parameter returns length of test cable.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_LINKSTATE

ViStatus Chr1102x_Q_LINKSTATE (ViSession instrumentHandle,

ViPBoolean LINK1320State);

Purpose

It queries LINK1320 ON/OFF.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LINK1320State

Variable Type       ViBoolean (passed by reference)

This parameter returns LINK1320 ON/OFF.

1:On

0:Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_measurementspeed

ViStatus Chr1102x_Q_measurementspeed (ViSession instrumentHandle,

ViPReal64 measureSpeed);

Purpose

It queries measuring speed.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

measureSpeed

Variable Type       ViReal64 (passed by reference)

This parameter returns the measurement speed.

Fast:0.025

Medium:0.065

Slow:0.5

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Measureparallelmode

ViStatus Chr1102x_Q_Measureparallelmode (ViSession instrumentHandle,

ViChar _VI_FAR measureParallelMode[]);

Purpose

It queries circuit measurement parallel mode.

Note:Only for 11025

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

measureParallelMode

Variable Type       ViChar[]

This parameter returns circuit measurement parallel mode.

When SENSe:FUNCtion:CONCurrent is OFF:

"FIMPedance"    Impedance measurement (equivalent series

circuit),"Set Measure Transformer State" will be set "off"

"FADMittance"   Admittance measurement (equivalent parallel

circuit),"Set Measure Transformer State" will be set "off"

When SENSe:FUNCtion:CONCurrent is ON :

'FIMPedance','FRESistance'  DCR measurement (equivalent series

circuit),"Set Measure Transformer State" will be set "ON"

'FADMittance','FRESistance' DCR measurement (equivalent parallel

circuit),"Set Measure Transformer State" will be set "ON"

'VOLTage:AC','IMPedance'    Transformer turn ratio measurement ,"Set

Measure Transformer State" will be set "ON"

'IMPedance','FADMittance'   Transformer mutual inductance ratio

measurement,"Set Measure Transformer State" will be set "ON"

'IMPedance','RESistance'    Transformer DC resistance

measurement,"Set Measure Transformer State" will be set "ON"

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Measurerange

ViStatus Chr1102x_Q_Measurerange (ViSession instrumentHandle,

ViPReal64 measureRange);

Purpose

It queries the measurement range.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

measureRange

Variable Type       ViReal64 (passed by reference)

This parameter retruns the measurement file.

Unit:Ohm

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Meatranformerstate

ViStatus Chr1102x_Q_Meatranformerstate (ViSession instrumentHandle,

ViPBoolean measureTransformerState);

Purpose

It queries whether to measure transformer parameter (DCR, L2, N, M and

R2) or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

measureTransformerState

Variable Type       ViBoolean (passed by reference)

This parameter returns whether to measure transformer parameter (DCR,

L2, N, M and R2) or not.

1:ON

0:Off

ON allows two commands of <SENSOR_FUNCTION>, it means to

perform transformer measurement.

OFF allows one command of <SENSOR_FUNCTION>, it means to

perform common measurement.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_NumRepMetAct

ViStatus Chr1102x_Q_NumRepMetAct (ViSession instrumentHandle,

ViPBoolean numRepresentMethodAction,

ViBoolean calculateNumber);

Purpose

It queries definition operation management is action or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

numRepresentMethodAction

Variable Type       ViBoolean (passed by reference)

This parameter returns definition operation management is action or

not.

Description:

ON (1)    :  Operation management is action.

OFF (0)   :  Operation management is not action.

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_NumRepMeth

ViStatus Chr1102x_Q_NumRepMeth (ViSession instrumentHandle,

ViChar _VI_FAR numericalRepresentMethod[],

ViBoolean calculateNumber);

Purpose

It queries numerical represent method.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

numericalRepresentMethod

Variable Type       ViChar[]

This parameter returns Numerical Represent Method

Ture(1):PCNT (Percentage)

False(0):DEV (Bias absolute value)

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_OPC

ViStatus Chr1102x_Q_OPC (ViSession instrumentHandle,

ViPInt32 OPCStatus);

Purpose

When completed all operation, notify the 11022/11025 set ASCII character

1 to output queue.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

OPCStatus

Variable Type       ViInt32 (passed by reference)

This parameter returns OPC Status

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Oupputimpedance

ViStatus Chr1102x_Q_Oupputimpedance (ViSession instrumentHandle,

ViChar _VI_FAR outputImpedance[]);

Purpose

It queries output impedance mode

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

outputImpedance

Variable Type       ViChar[]

This parameter returns output impedance mode

Unit: OHM

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Para1

ViStatus Chr1102x_Q_Para1 (ViSession instrumentHandle,

ViChar _VI_FAR para1OfMeasurement[]);

Purpose

It queries the current measurement master parameter.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

para1OfMeasurement

Variable Type       ViChar[]

This parameter returns the current measurement master parameter.

Return Parameter:

{REAL|MLINear|CP|CS|LP|LS|NONE}

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Para2

ViStatus Chr1102x_Q_Para2 (ViSession instrumentHandle,

ViChar _VI_FAR para2OfMeasurement[]);

Purpose

It queries the current measurement secondary parameter.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

para2OfMeasurement

Variable Type       ViChar[]

This parameter returns the current measurement secondary parameter.

Return Parameter:

{PHASe|D|Q|REAL|MLINear|LP|INV}

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_PRI_HI_LIMIT

ViStatus Chr1102x_Q_PRI_HI_LIMIT (ViSession instrumentHandle,

ViPReal64 PRI_HI_LIMITOfCompare,

ViBoolean calculateNumber);

Purpose

It queries the high limit value.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

PRI_HI_LIMITOfCompare

Variable Type       ViReal64 (passed by reference)

This parameter retruns the low limit value.

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_PRI_LO

ViStatus Chr1102x_Q_PRI_LO (ViSession instrumentHandle,

ViPReal64 PRI_LO_LIMITOfCompare,

ViBoolean calculateNumber);

Purpose

It queries the low limit value.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

PRI_LO_LIMITOfCompare

Variable Type       ViReal64 (passed by reference)

This parameter retruns the low limit value.

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_SRE

ViStatus Chr1102x_Q_SRE (ViSession instrumentHandle, ViPInt32 SREValue);

Purpose

It queries the content of the Status Byte Enable Register

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

SREValue

Variable Type       ViInt32 (passed by reference)

This parameter returns SRE Value

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_STB

ViStatus Chr1102x_Q_STB (ViSession instrumentHandle,

ViPInt32 STBStatus);

Purpose

It queries the contents of the Status Bit Register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

STBStatus

Variable Type       ViInt32 (passed by reference)

This parameter returns STB Status

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Systemintegration

ViStatus Chr1102x_Q_Systemintegration (ViSession instrumentHandle,

ViPInt32 systemIntegration);

Purpose

It queries calculates period selection function for deciding every

sampling period when quick measuring.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

systemIntegration

Variable Type       ViInt32 (passed by reference)

This parameter returns calculates period selection function for

deciding every sampling period when quick measuring.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Testfrequency

ViStatus Chr1102x_Q_Testfrequency (ViSession instrumentHandle,

ViPReal64 testFrequency);

Purpose

It queries the test frequency now.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

testFrequency

Variable Type       ViReal64 (passed by reference)

This parameter returns the test frequency now.

Unit:Hz

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Testvoltage

ViStatus Chr1102x_Q_Testvoltage (ViSession instrumentHandle,

ViPReal64 testVoltage);

Purpose

It queries test voltage now.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

testVoltage

Variable Type       ViReal64 (passed by reference)

This parameter returns test voltage now.

Unit: V

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_TriggerDelayTime

ViStatus Chr1102x_Q_TriggerDelayTime (ViSession instrumentHandle,

ViPReal64 triggerDealyTime);

Purpose

It queries trigger delay time.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

triggerDealyTime

Variable Type       ViReal64 (passed by reference)

This parameter returns trigger delay time.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_Triggeredge

ViStatus Chr1102x_Q_Triggeredge (ViSession instrumentHandle,

ViChar _VI_FAR triggerEdge[]);

Purpose

It queries TRIGGER mode (RISING or FALLING trigger.)

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

triggerEdge

Variable Type       ViChar[]

This parameter returns TRIGGER mode (RISING or FALLING trigger.)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_TriggerSource

ViStatus Chr1102x_Q_TriggerSource (ViSession instrumentHandle,

ViChar _VI_FAR triggerSource[]);

Purpose

It queries trigger mode now.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

triggerSource

Variable Type       ViChar[]

This parameter returns trigger mode now.

Description:

BUS      :BUS Trigger

EXTernal :External Trigger

INTernal :Internal Trigger

MANual   :Manual Trigger

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_TST

ViStatus Chr1102x_Q_TST (ViSession instrumentHandle,

ViChar _VI_FAR TSTValue[]);

Purpose

It performs self-test and return test result of error total value. If

there is no error the 11022/11025 return a 0.

RAM1               1

RAM2               2

EEPROM             4

FPGA               8

Calibration Data   16

Main Board Test    32

Key Control        64

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

TSTValue

Variable Type       ViChar[]

This parameter retruns TST Value

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_VIDisplay

ViStatus Chr1102x_Q_VIDisplay (ViSession instrumentHandle,

ViPBoolean vm_ImDisplayState,

ViBoolean calculateNumber);

Purpose

It queries whether turn on Vm/Im function.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

vm_ImDisplayState

Variable Type       ViBoolean (passed by reference)

This parameter returns whether turn on Vm/Im function.

Description:

ON (1) :Turn on

OFF (0) :Turn off

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Set current (Im):0

Set voltage (Vm):1

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Q_VoltageOffsetState

ViStatus Chr1102x_Q_VoltageOffsetState (ViSession instrumentHandle,

ViPBoolean voltageOffsetState);

Purpose

It queries if BIAS V function is enabled.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

voltageOffsetState

Variable Type       ViBoolean (passed by reference)

This parameter returns if BIAS V function is enabled.

1:On

0:Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueriesID_String

ViStatus Chr1102x_QueriesID_String (ViSession instrumentHandle,

ViChar _VI_FAR queriesIDString[]);

Purpose

It queries identification string which consists of four fields.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

queriesIDString

Variable Type       ViChar[]

It queries identification string which consists of four fields.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_Query_BIN9_State

ViStatus Chr1102x_Query_BIN9_State (ViSession instrumentHandle,

ViPBoolean BINMore9State);

Purpose

It queries BINNING >9 state is open or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BINMore9State

Variable Type       ViBoolean (passed by reference)

This parameter retruns It queries BINNING >9 state is open or not.

1: ON (open)

0: OFF (close)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryAUX_HI

ViStatus Chr1102x_QueryAUX_HI (ViSession instrumentHandle,

ViPReal64 AUX_HI);

Purpose

It queries the upper value of secondary parameter B

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

AUX_HI

Variable Type       ViReal64 (passed by reference)

This parameter retruns the upper value of secondary parameter B

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryAUX_LO

ViStatus Chr1102x_QueryAUX_LO (ViSession instrumentHandle,

ViPReal64 AUX_LO);

Purpose

It queries the lower value of secondary parameter B

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

AUX_LO

Variable Type       ViReal64 (passed by reference)

This parameter retruns the lower value of secondary parameter B

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryBias

ViStatus Chr1102x_QueryBias (ViSession instrumentHandle,

ViChar _VI_FAR biasDEVORPCNT[],

ViBoolean calculateNumber);

Purpose

It queries the parameter which can use with Num Represent Method

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

biasDEVORPCNT

Variable Type       ViChar[]

This parameter returns Bias absolute value (DEV) or percentage

(PCNT).

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryBinningMode

ViStatus Chr1102x_QueryBinningMode (ViSession instrumentHandle,

ViChar _VI_FAR binningMode[]);

Purpose

It queries Bining Mode

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binningMode

Variable Type       ViChar[]

This parameter returns Bining Mode (ABS/PCNT)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryBINNumber_9

ViStatus Chr1102x_QueryBINNumber_9 (ViSession instrumentHandle,

int *BIN_Num_9);

Purpose

It queries BIN NO when BIN NO>9 function in SYSTEM CONFIG is enabled.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BIN_Num_9

Variable Type       int (passed by reference)

This parameter returns BIN NO when BIN NO>9 function in SYSTEM CONFIG

is enabled.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryCorrectionValue

ViStatus Chr1102x_QueryCorrectionValue (ViSession instrumentHandle,

ViInt32 correctionProcedure,

ViPReal64 correctionValue1,

ViPReal64 correctionValue2);

Purpose

It queries correction numeral.

OPEN correction value in the current range, that is G, B.

SHORT correction value in the current range, that is R, X.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionProcedure

Variable Type       ViInt32

It selects query correction numeral.

1:It is OPEN correction numeral, that is G, B.

2:It is SHORT correction numeral, that is R, X.

correctionValue1

Variable Type       ViReal64 (passed by reference)

This parameter returns correction numeral.

When is OPEN correction numeral, that is G

When is SHORT correction numeral, that is R

correctionValue2

Variable Type       ViReal64 (passed by reference)

This parameter returns correction numeral.

When is OPEN correction numeral, that is B.

When is SHORT correction numeral, that is X.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryDataValueofCompare

ViStatus Chr1102x_QueryDataValueofCompare (ViSession instrumentHandle,

ViInt32 selectParameter,

ViPReal64 nominalValueOfCompare);

Purpose

It queries REF1 (the nominal value of COMPARE in primary parameter),REF2

(the nominal value of COMPARE in secondary parameter),IMON (measuring

current value),VMON (measuring voltage value).

When Im/Vm Display function is OFF, return values of IMON, VMON is 0.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

selectParameter

Variable Type       ViInt32

Selects query REF1 (the nominal value of COMPARE in primary

parameter),REF2 (the nominal value of COMPARE in secondary

parameter),IMON (measuring current value),VMON (measuring voltage

value).

REF1:0

REF2:1

IMON:2

VMON:3

nominalValueOfCompare

Variable Type       ViReal64 (passed by reference)

It parameter returns REF1 (the nominal value of COMPARE in primary

parameter),REF2 (the nominal value of COMPARE in secondary

parameter),IMON (measuring current value),VMON (measuring voltage

value).

When Im/Vm Display function is OFF, return values of IMON, VMON

is 0.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryLowLimitBin

ViStatus Chr1102x_QueryLowLimitBin (ViSession instrumentHandle,

ViInt32 binNumber,

ViPReal64 lowLinit);

Purpose

It queries the lower limit value of each BIN

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binNumber

Variable Type       ViInt32

It sets the number of BINs.

Valid Range: 1 ~ 8.

lowLinit

Variable Type       ViReal64 (passed by reference)

This parameter retruns the lower limit value of each BIN

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryNominal

ViStatus Chr1102x_QueryNominal (ViSession instrumentHandle,

ViPReal64 nominal);

Purpose

It queries the nominal value of primary parameter A

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

nominal

Variable Type       ViReal64 (passed by reference)

This parameter retruns the nominal value of primary parameter A

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryUpperLinitofBin

ViStatus Chr1102x_QueryUpperLinitofBin (ViSession instrumentHandle,

ViInt32 binNumber,

ViPReal64 upperLinit);

Purpose

It queries the upper limit value of each BINs

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binNumber

Variable Type       ViInt32

It sets the number of BINs.

Valid Range: 1 ~ 8.

upperLinit

Variable Type       ViReal64 (passed by reference)

This parameter retruns the upper limit value of each BINs

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueryVersion

ViStatus Chr1102x_QueryVersion (ViSession instrumentHandle,

ViChar _VI_FAR queryVersion[]);

Purpose

The SCPI version corresponding value is Y.V, Y means Year, V means

version.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

queryVersion

Variable Type       ViChar[]

The SCPI version corresponding value is Y.V, Y means Year, V means

version.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueyStatusOperation

ViStatus Chr1102x_QueyStatusOperation (ViSession instrumentHandle,

ViPInt32 queyStatusOperation);

Purpose

It returns event register status of operation status group, and clear it

after read event register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

queyStatusOperation

Variable Type       ViInt32 (passed by reference)

It returns event register status of operation status group, and clear

it after read event register.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_QueyStatusOperationCondition

ViStatus Chr1102x_QueyStatusOperationCondition

(ViSession instrumentHandle,

ViPInt32 queyStatusOperationCondition);

Purpose

It returns condition register status of operation condition group, and

don't clear it after read condition register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

queyStatusOperationCondition

Variable Type       ViInt32 (passed by reference)

It returns condition register status of operation condition group,

and don't clear it after read condition register.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_ReadInstrData

ViStatus Chr1102x_ReadInstrData (ViSession instrumentHandle,

ViInt32 numberOfBytesToRead,

ViChar _VI_FAR readBuffer[],

ViPInt32 numBytesRead);

Purpose

This function reads data from the instrument.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

numberOfBytesToRead

Variable Type       ViInt32

Pass the maximum number of bytes to read from the instruments.

Valid Range:  0 to the number of elements in the Read Buffer.

Default:  0

readBuffer

Variable Type       ViChar[]

After this function executes, this parameter contains the data that

was read from the instrument.

numBytesRead

Variable Type       ViInt32 (passed by reference)

Returns the number of bytes actually read from the instrument and

stored in the Read Buffer.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_reset

ViStatus Chr1102x_reset (ViSession instrumentHandle);

Purpose

This function resets the instrument to a known state and sends

initialization commands to the instrument.  The initialization commands

set instrument settings such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0102  Reset not supported.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_revision_query

ViStatus Chr1102x_revision_query (ViSession instrumentHandle,

ViChar _VI_FAR instrumentDriverRevision[],

ViChar _VI_FAR firmwareRevision[]);

Purpose

This function returns the revision numbers of the instrument driver and

instrument firmware.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

instrumentDriverRevision

Variable Type       ViChar[]

Returns the instrument driver software revision numbers in the form

of a string.

You must pass a ViChar array with at least 256 bytes.

firmwareRevision

Variable Type       ViChar[]

Returns the instrument firmware revision numbers in the form of a

string.

You must pass a ViChar array with at least 256 bytes.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0105  Revision Query not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_self_test

ViStatus Chr1102x_self_test (ViSession instrumentHandle,

ViPInt16 selfTestResult,

ViChar _VI_FAR selfTestMessage[]);

Purpose

This function runs the instrument's self test routine and returns the

test result(s).

It performs self-test and return test result of error total value.  If

there is no error the 11025 return a 0.

RAM1                 1

RAM2                 2

EEPROM               4

FPGA                 8

Calibration Data    16

Main Board Test     32

Key Control         64

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

selfTestResult

Variable Type       ViInt16 (passed by reference)

This control contains the value returned from the instrument self

test.  Zero means success.  For any other code, see the device's

operator's manual.

Self-Test Code    Description

---------------------------------------

0              Passed self test

1              Self test failed

selfTestMessage

Variable Type       ViChar[]

Returns the self-test response string from the instrument. See the

device's operation manual for an explanation of the string's

contents.

You must pass a ViChar array with at least 256 bytes.

It performs self-test and return test result of error total value.

If there is no error the 11025 return a 0.

RAM1                 1

RAM2                 2

EEPROM               4

FPGA                 8

Calibration Data    16

Main Board Test     32

Key Control         64

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetAbort

ViStatus Chr1102x_SetAbort (ViSession instrumentHandle);

Purpose

It terminates processing trigger system immediately, that is say to reset

trigger system.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetAttributeViBoolean

ViStatus Chr1102x_SetAttributeViBoolean (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViBoolean attributeValue);

Purpose

This function sets the value of a ViBoolean attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViBoolean

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViBoolean are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViBoolean

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_SetAttributeViInt32

ViStatus Chr1102x_SetAttributeViInt32 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViInt32 attributeValue);

Purpose

This function sets the value of a ViInt32 attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViInt32 type.

If you choose to see all IVI attributes, the data types appear

to the right of the attribute names in the list box.

Attributes with data types other than ViInt32 are dim.  If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViInt32

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_SetAttributeViReal64

ViStatus Chr1102x_SetAttributeViReal64 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViReal64 attributeValue);

Purpose

This function sets the value of a ViReal64 attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViReal64

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViReal64 are dim.  If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViReal64

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_SetAttributeViSession

ViStatus Chr1102x_SetAttributeViSession (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViSession attributeValue);

Purpose

This function sets the value of a ViSession attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViSession

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViSession are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViSession

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_SetAttributeViString

ViStatus Chr1102x_SetAttributeViString (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViChar _VI_FAR attributeValue[]);

Purpose

This function sets the value of a ViString attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViString

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViString are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViChar[]

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFA0001  Instrument error. Call Chr1102x_error_query.

BFFA000C  Invalid attribute.

BFFA000D  Attribute is not writable.

BFFA000E  Attribute is not readable.

BFFA000F  Invalid parameter.

BFFA0010  Invalid value.

BFFA0012  Attribute not supported.

BFFA0013  Value not supported.

BFFA0014  Invalid type.

BFFA0015  Types do not match.

BFFA0016  Attribute already has a value waiting to be updated.

BFFA0018  Not a valid configuration.

BFFA0019  Requested item does not exist or value not available.

BFFA001A  Requested attribute value not known.

BFFA001B  No range table.

BFFA001C  Range table is invalid.

BFFA001F  No channel table has been built for the session.

BFFA0020  Channel name specified is not valid.

BFFA0044  Channel name required.

BFFA0045  Channel name not allowed.

BFFA0046  Attribute not valid for channel.

BFFA0047  Attribute must be channel based.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

3FFF0085  The status value you passed is unknown.

________________________________________________________________________________

Chr1102x_SetAutoRangeState

ViStatus Chr1102x_SetAutoRangeState (ViSession instrumentHandle,

ViBoolean autoRangeState);

Purpose

It sets whether turn on auto file selection mode.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

autoRangeState

Variable Type       ViBoolean

It sets whether turn on auto file selection mode.

Ture(1):On

False(0):Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetAUX_HIofBin

ViStatus Chr1102x_SetAUX_HIofBin (ViSession instrumentHandle,

ViReal64 setAUX_HIOfBIN);

Purpose

It sets AUX-HI value of BINNING, that is the upper limit value of

measurement parameter B.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setAUX_HIOfBIN

Variable Type       ViReal64

It sets AUX-HI value of BINNING, that is the upper limit value of

measurement parameter B.

Vaild Range: -999900000000000 ~ 999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetAUX_LOofBin

ViStatus Chr1102x_SetAUX_LOofBin (ViSession instrumentHandle,

ViReal64 setAUX_LOOfBIN);

Purpose

It sets AUX-LO value of BINNING, that is the lower limit value of

measurement parameter B.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setAUX_LOOfBIN

Variable Type       ViReal64

It sets AUX-LO value of BINNING, that is the lower limit value of

measurement parameter B.

Vaild Range: -999900000000000 ~ 999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetAverageCount

ViStatus Chr1102x_SetAverageCount (ViSession instrumentHandle,

ViInt32 setAverageCount);

Purpose

It sets average time.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setAverageCount

Variable Type       ViInt32

This parameter sets average time

Range:1~256

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBacktoDefaultState

ViStatus Chr1102x_SetBacktoDefaultState (ViSession instrumentHandle);

Purpose

It sets 11025 back to default status, but do not change key lock status.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBeeper

ViStatus Chr1102x_SetBeeper (ViSession instrumentHandle);

Purpose

Buzzer sound beep immediately.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBeeperState

ViStatus Chr1102x_SetBeeperState (ViSession instrumentHandle,

ViInt32 beeperState);

Purpose

It sets whether buzzer is turned on or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

beeperState

Variable Type       ViInt32

It sets whether buzzer is turned on or not.

0:OFF - turn off buzzer

1:ON - turn on buzzer (small)

2:LARGE - turn on buzzer (large)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBIASCurrentState

ViStatus Chr1102x_SetBIASCurrentState (ViSession instrumentHandle,

ViBoolean setBIASCurrentState);

Purpose

It sets the output condition of bias current source.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setBIASCurrentState

Variable Type       ViBoolean

It sets the output condition of bias current source.

Ture(1):On

False(0):Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBIASCurrentValue

ViStatus Chr1102x_SetBIASCurrentValue (ViSession instrumentHandle,

ViReal64 BIASCurernt);

Purpose

It sets bias current value.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BIASCurernt

Variable Type       ViReal64

It returns and sets bias current value.

Be able to set the range of bias current value of a 1320 Bias Current

Source 0-10A as 0.000A~10.000A.  Be able to set the range of bias

current value of a 1320 Bias Current Source 0-20A as 0.000A~20.000A.

When a 1320 Bias Current Source 0-20A connects a 1320S, bias current

value range can be set is 0.000A ~ 40.000A.  A 1320 Bias Current

Source 0-20A can connect four 1320S maximumly, thus bias current

value can be set is 0.000A ~ 100.000A.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBINHI9

ViStatus Chr1102x_SetBINHI9 (ViReal64 HIValue,

ViSession instrumentHandle);

Purpose

It sets the upper limit value of BIN NO when BIN NO>9 function in SYSTEM

CONFIG is enabled.

When BINning MODE is ABS, it is the absolute value of inputted upper

limit.  When BINning MODE is PCNT, it is the percentage of inputted upper

limit.

Parameter List

HIValue

Variable Type       ViReal64

It sets BIN high limit value when BIN NO greater than 9.

Vaild Range:

ABS MODE: -999900000000000 ~ 999900000000000

PCNT MODE: -999.99 ~ 999.99

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBINLo9

ViStatus Chr1102x_SetBINLo9 (ViReal64 LOValue,

ViSession instrumentHandle);

Purpose

It sets the lower limit value of BIN NO when BIN NO>9 function in SYSTEM

CONFIG is enabled.

When BINning MODE is ABS, it is the absolute value of inputted upper

limit.  When BINning MODE is PCNT, it is the percentage of inputted upper

limit.

Parameter List

LOValue

Variable Type       ViReal64

It sets BIN low limit value when BIN NO greater than 9.

Vaild Range:

ABS MODE: -999900000000000 ~ 999900000000000

PCNT MODE: -999.99 ~ 999.99

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBINMore9State

ViStatus Chr1102x_SetBINMore9State (ViSession instrumentHandle,

ViBoolean BINMore9State);

Purpose

It sets BINNING >9 state is open or not.

Ture(1):ON

False(0):OFF

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BINMore9State

Variable Type       ViBoolean

It sets BINNING >9 state is open or not.

On:1

Off:0

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBinningMode

ViStatus Chr1102x_SetBinningMode (ViSession instrumentHandle,

ViBoolean binningMode);

Purpose

It sets MODE to absolute value or percentage.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binningMode

Variable Type       ViBoolean

It sets MODE to absolute value or percentage.

Ture(1):Absolute

False(0):Percentage

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBinningState

ViStatus Chr1102x_SetBinningState (ViSession instrumentHandle,

ViBoolean binningState);

Purpose

It sets BINNING is open or not.

Ture(1):Open

False(0):Close

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binningState

Variable Type       ViBoolean

It sets BINNING is open or not.

Ture(1):Open

False(0):Close

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBINNumber_9

ViStatus Chr1102x_SetBINNumber_9 (ViSession instrumentHandle,

ViInt32 BINNo);

Purpose

It sets BIN NO when BIN NO>9 function in SYSTEM CONFIG is enabled.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

BINNo

Variable Type       ViInt32

It sets maximum number value when BIN NO greater than 9.

Range :1~99

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBuzzerActionType

ViStatus Chr1102x_SetBuzzerActionType (ViSession instrumentHandle,

ViBoolean buzzerActionType);

Purpose

It selects buzzer action is PULSE or CONTINUOUS when BINSORTING and

COMPARE.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

buzzerActionType

Variable Type       ViBoolean

It selects buzzer action is PULSE or CONTINUOUS when BINSORTING and

COMPARE.

Ture(1):CONTINUOUS

False(0):PULSE

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBuzzerOutputAction

ViStatus Chr1102x_SetBuzzerOutputAction (ViSession instrumentHandle,

ViBoolean calculateNumber,

ViBoolean buzzerOutputEnable);

Purpose

It sets whether buzzer is action or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

buzzerOutputEnable

Variable Type       ViBoolean

It sets whether buzzer is action or not.

Description:

ON (1):    Turn on buzzer.

OFF (0):   Turn off buzzer.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetBuzzerOutputDefine

ViStatus Chr1102x_SetBuzzerOutputDefine (ViSession instrumentHandle,

ViBoolean calculateNumber,

ViBoolean buzzerOutputDefine);

Purpose

Buzzer output definition.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

buzzerOutputDefine

Variable Type       ViBoolean

Buzzer output definition.

Description:

FAIL(0):When the comparison result is FAIL, buzzer will be action.

PASS(1):When the comparison result is PASS, buzzer will be action.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCalc1ofDualFirstMeasure

ViStatus Chr1102x_SetCalc1ofDualFirstMeasure (ViSession instrumentHandle,

ViInt32 calc1OfDualFirstMeasure);

Purpose

It sets master parameter of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calc1OfDualFirstMeasure

Variable Type       ViInt32

It sets master parameter of first measure of the dual function.

0:REAL

1:MLINEAR

2:CP

3:CS

4:LP

5:LS

6:ZS

7:RS

8:RP

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCalc2ofDual2ndMeasure

ViStatus Chr1102x_SetCalc2ofDual2ndMeasure (ViSession instrumentHandle,

ViInt32 calc2OfDual2ndMeasure);

Purpose

It sets secondary parameter of secondary measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calc2OfDual2ndMeasure

Variable Type       ViInt32

It sets secondary parameter of secondary measure of the dual

function.

0:Imaginary

1:Phase

2:D

3:Q

4:REAL

5:RS

6:XS

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCalc2ofDualFirstMeasure

ViStatus Chr1102x_SetCalc2ofDualFirstMeasure (ViSession instrumentHandle,

ViInt32 calc2OfDualFirstMeasure);

Purpose

It sets secondary parameter of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calc2OfDualFirstMeasure

Variable Type       ViInt32

It sets secondary parameter of first measure of the dual function.

0:Imaginary

1:Phase

2:D

3:Q

4:REAL

5:RS

6:XS

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCompareLimitAction

ViStatus Chr1102x_SetCompareLimitAction (ViSession instrumentHandle,

ViBoolean calculateNumber,

ViBoolean compareLimitAction);

Purpose

It sets whether the comparison function is action or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

compareLimitAction

Variable Type       ViBoolean

It sets whether the comparison function is action or not.

Description:

ON(1)     Turn on comparison function.

OFF(0)    Turn off comparison function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCorr_DataShort100MOhm

ViStatus Chr1102x_SetCorr_DataShort100MOhm (ViSession instrumentHandle,

ViReal64 correctionValueR,

ViReal64 correctionValueX);

Purpose

It sets SHORT correction value of RANGE 100m[, that is R, X

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionValueR

Variable Type       ViReal64

It is R of SHORT correction value of RANGE 100m[

Range:-999900000000000~999900000000000

correctionValueX

Variable Type       ViReal64

It is X of SHORT correction value of RANGE 100m[

Range:-999900000000000~999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCorrectionDataOfOpen

ViStatus Chr1102x_SetCorrectionDataOfOpen (ViSession instrumentHandle,

ViReal64 correctionValueG,

ViReal64 correctionValueB);

Purpose

It sets or queries OPEN correction value, that is G, B.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionValueG

Variable Type       ViReal64

It is G of OPEN correction numeral.

Range:-999900000000000~999900000000000

correctionValueB

Variable Type       ViReal64

It is B of OPEN correction numeral.

Range:-999900000000000~999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCorrectionDataOfShort

ViStatus Chr1102x_SetCorrectionDataOfShort (ViSession instrumentHandle,

ViReal64 correctionValueR,

ViReal64 correctionValueX);

Purpose

It sets or queries correction value of SHORT>=10,that is R, X

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionValueR

Variable Type       ViReal64

It is R of Short value>=10 ohm correction numeral.

Range:-999900000000000~999900000000000

correctionValueX

Variable Type       ViReal64

It is X of Short value>=10 ohm correction numeral.

Range:-999900000000000~999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCorrectionDataOfShort1Ohm

ViStatus Chr1102x_SetCorrectionDataOfShort1Ohm

(ViSession instrumentHandle, ViReal64 correctionValueR,

ViReal64 correctionValueX);

Purpose

It sets SHORT correction value of RANGE 1[, that is R, X

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionValueR

Variable Type       ViReal64

It is correction numeral.

It is R of SHORT correction value of RANGE 1[

Range:-999900000000000~999900000000000

correctionValueX

Variable Type       ViReal64

It is X of SHORT correction value of RANGE 1[

Range:-999900000000000~999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetCorrectionProcedure

ViStatus Chr1102x_SetCorrectionProcedure (ViSession instrumentHandle,

ViInt32 correctionProcedure);

Purpose

It performs OPEN, SHORT correction procedure.

NOTE:

"DCR Short" only for 11025

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

correctionProcedure

Variable Type       ViInt32

It performs OPEN, SHORT correction procedure.

Description:

1.Multiple Frequency Open: It performs Multiple Frequency OPEN

correction procedure.

2.Multiple Frequency Short: It performs Multiple Frequency Short

correction procedure.

3.Single Frequency Open: It performs Single Frequency OPEN correction

procedure.

4.Single Frequency Short:It performs Single Frequency Short

correction procedure.

5.DCR Short: It performs DCR Short correction procedure.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualAutoofFirstMeasure

ViStatus Chr1102x_SetDualAutoofFirstMeasure (ViSession instrumentHandle,

ViBoolean dualAutoOfFirstMeasure);

Purpose

It sets auto state of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualAutoOfFirstMeasure

Variable Type       ViBoolean

It sets auto state of first measure of the dual function.

Ture(1):ON

False(0):OFF

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualAutoofSecondMeasure

ViStatus Chr1102x_SetDualAutoofSecondMeasure (ViSession instrumentHandle,

ViBoolean dualAutoOfSecondMeasure);

Purpose

It sets auto state of second measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualAutoOfSecondMeasure

Variable Type       ViBoolean

It sets auto state of second measure of the dual function.

Ture(1):ON

False(0)0:OFF

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualBinningMode

ViStatus Chr1102x_SetDualBinningMode (ViSession instrumentHandle,

ViBoolean setDualBinningMode);

Purpose

It sets binning mode of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualBinningMode

Variable Type       ViBoolean

It sets binning mode of the dual function.

False(0):ABS

Ture(1):PERCENT

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualFirstBinningAux_Hi

ViStatus Chr1102x_SetDualFirstBinningAux_Hi (ViSession instrumentHandle,

ViReal64 setDualFirstBinningAux_Hi);

Purpose

It sets first binning aux-Hi value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualFirstBinningAux_Hi

Variable Type       ViReal64

It sets first binning aux-Hi value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualFirstBinningAux_Lo

ViStatus Chr1102x_SetDualFirstBinningAux_Lo (ViSession instrumentHandle,

ViReal64 setDualFirstBinningAux_Lo);

Purpose

It sets first binning aux-Lo value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualFirstBinningAux_Lo

Variable Type       ViReal64

It sets first binning aux-Lo value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualFirstBinningNominal

ViStatus Chr1102x_SetDualFirstBinningNominal (ViSession instrumentHandle,

ViReal64 setDualFirstBinningNominal);

Purpose

It sets first binning nominal value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualFirstBinningNominal

Variable Type       ViReal64

It sets first binning nominal value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualFirstMeasureRange

ViStatus Chr1102x_SetDualFirstMeasureRange (ViSession instrumentHandle,

ViInt32 dualFirstMeasureRange);

Purpose

It sets range value of first measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualFirstMeasureRange

Variable Type       ViInt32

It sets range value of first measure of the dual function.

0:0.1 ohm

1:1 ohm

2:10 ohm

3:100 ohm

4:1k ohm

5:10k ohm

6:100k ohm

7:1M ohm

8:UP key simulation

9:Down key simulation

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualFirstSourceFreq

ViStatus Chr1102x_SetDualFirstSourceFreq (ViSession instrumentHandle,

ViInt32 setDualFirstSourceFreq);

Purpose

It sets source frequnecy range of first test of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualFirstSourceFreq

Variable Type       ViInt32

It sets frequnecy range of first measure of the dual function.

50:50 Hz

60:60 Hz

100:100 Hz

120:120 Hz

1000:1K Hz

10000:10K Hz

20000:20K Hz

40000:40K Hz

50000:50K Hz

100000:100K Hz

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualSecondBinningAux_Hi

ViStatus Chr1102x_SetDualSecondBinningAux_Hi (ViSession instrumentHandle,

ViReal64 setDualSecondBinningAux_Hi);

Purpose

It sets second binning aux-Hi value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualSecondBinningAux_Hi

Variable Type       ViReal64

It sets second binning aux-Hi value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualSecondBinningAux_Lo

ViStatus Chr1102x_SetDualSecondBinningAux_Lo (ViSession instrumentHandle,

ViReal64 setDualSecondBinningAux_Lo);

Purpose

It sets second binning aux-Lo value of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualSecondBinningAux_Lo

Variable Type       ViReal64

It sets second binning aux-Lo value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualSecondMeasureRange

ViStatus Chr1102x_SetDualSecondMeasureRange (ViSession instrumentHandle,

ViInt32 dualSecondMeasureRange);

Purpose

It sets range value of second measure of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

dualSecondMeasureRange

Variable Type       ViInt32

It sets range value of second measure of the dual function.

0:0.1 ohm

1:1 ohm

2:10 ohm

3:100 ohm

4:1k ohm

5:10k ohm

6:100k ohm

7:1M ohm

8:UP key simulation

9:Down key simulation

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetDualSecondSourceFreq

ViStatus Chr1102x_SetDualSecondSourceFreq (ViSession instrumentHandle,

ViInt32 setDualSecondSourceFreq);

Purpose

It sets source frequnecy range of second test of the dual function.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setDualSecondSourceFreq

Variable Type       ViInt32

It sets frequnecy range of second measure of the dual function.

50:50 Hz

60:60 Hz

100:100 Hz

120:120 Hz

1000:1K Hz

10000:10K Hz

20000:20K Hz

40000:40K Hz

50000:50K Hz

100000:100K Hz

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetEnableRegisterStatus

ViStatus Chr1102x_SetEnableRegisterStatus (ViSession instrumentHandle,

ViInt32 setEnableRegisterStatus);

Purpose

It sets enable register status.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setEnableRegisterStatus

Variable Type       ViInt32

It sets enable register status.

Range:0~255

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetESE

ViStatus Chr1102x_SetESE (ViSession instrumentHandle, ViInt32 ESEValue);

Purpose

It sets the bits of the Standard Event Status Enable Register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

ESEValue

Variable Type       ViInt32

It sets the bits of the Standard Event Status Enable Register.

Valid Range:0 ~ 255.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetFirstBinningLower

ViStatus Chr1102x_SetFirstBinningLower (ViSession instrumentHandle,

ViInt32 binNumber,

ViReal64 firstBinningLower);

Purpose

It sets the lower limit value of each BINs.  When BINning: MODE ABS, it

is the absolute value of inputted upper limit.  When BINning: MODE PCNT,

it is the percentage of inputted lower limit.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binNumber

Variable Type       ViInt32

It sets the number of BINs of dual function.

Valid Range: 1 ~ 4.

firstBinningLower

Variable Type       ViReal64

It sets the upper limit value of each BINs of dual function.  When

BINning: MODE ABS, it is the absolute value of inputted upper limit.

When BINning: MODE PCNT, it is the percentage of inputted upper

limit.

Vaild Range:

PCNT:-999.9~999.9

ABS :-999900000000000~999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetFirstBinningUpper

ViStatus Chr1102x_SetFirstBinningUpper (ViSession instrumentHandle,

ViInt32 binNumber,

ViReal64 firstBinningUpper);

Purpose

It sets the upper limit value of each BINs.  When BINning: MODE ABS, it

is the absolute value of inputted upper limit.  When BINning: MODE PCNT,

it is the percentage of inputted upper limit.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binNumber

Variable Type       ViInt32

It sets the number of BINs of dual function.

Valid Range: 1 ~ 4.

firstBinningUpper

Variable Type       ViReal64

It sets the upper limit value of each BINs of dual function.  When

BINning: MODE ABS, it is the absolute value of inputted upper limit.

When BINning: MODE PCNT, it is the percentage of inputted upper

limit.

Vaild Range:

PCNT:-999.9~999.9

ABS :-999900000000000~999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetHandlerMode

ViStatus Chr1102x_SetHandlerMode (ViSession instrumentHandle,

ViBoolean handlerMode);

Purpose

It selects HANDLER MODE.The factory default value is CLEAR.

CLEAR : When using Handler interface, each time before measurement the

last test result will be cleared (Pass or Fail).

HOLD : When using Handler interface, test result (Pass or Fail) will

remain until next test result is displayed.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

handlerMode

Variable Type       ViBoolean

It selects HANDLER MODE.The factory default value is CLEAR.

Fasle (0):CLEAR : When using Handler interface, each time before

measurement the last test result will be cleared (Pass or Fail).

Ture (1):HOLD : When using Handler interface, test result (Pass or

Fail) will remain until next test result is displayed.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetInitiateContinuous

ViStatus Chr1102x_SetInitiateContinuous (ViSession instrumentHandle,

ViBoolean initiateContinuous);

Purpose

It sets whether initial trigger system sequentially.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

initiateContinuous

Variable Type       ViBoolean

It sets whether initial trigger system sequentially.

False(0) :OFF,No sequential initial trigger system

Ture(1) :ON, With sequential initial trigger system

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetInitiateImmediate

ViStatus Chr1102x_SetInitiateImmediate (ViSession instrumentHandle);

Purpose

All sequences escape from idle status and enter initial status.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetItemofMeasurement

ViStatus Chr1102x_SetItemofMeasurement (ViSession instrumentHandle,

ViInt32 itemOfMeasurement);

Purpose

It sets the current measurement parameter.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

itemOfMeasurement

Variable Type       ViInt32

It sets the current measurement parameter as following description.

0:Zs-Phase :Impedance absolute value and Phase.

1:Rs-Xs:Impedance real and impedance imaginary.

2:Cp-D: Equivalent parallel capacity and dissipation factor.

3:Cp-Q: Equivalent parallel capacity and quality factor.

4:Cs-D: Equivalent series capacity and dissipation factor.

5:Cs-Q: Equivalent series capacity and quality factor.

6:Cs-Rs: Equivalent series capacity and equivalent parallel

resistance.

7:Lp-D:  Equivalent parallel induction and dissipation factor.

8:Lp-Q: Equivalent parallel induction and quality factor.

9:Lp-Rs: Equivalent parallel induction and equivalent parallel

resistance.

10:Ls-Rdc: Equivalent series induction and impedance real.

11:Lp-Rdc: Equivalent parallel induction and impedance real.

12:L2a-N:Transformer turn ratio measurement for 11025 model use only.

(L2A parameter is being set.)

13:L2b-N:Transformer turn ratio measurement for 11025 model use only.

(L2B parameter is being set.)

14:L2a-1/N: The inverse of transformer turn ratio measurement for

11025 model use only. (L2A parameter is being set.)

15:L2b-1/N: The inverse of transformer turn ratio measurement for

11025 model use only. (L2B parameter is being set.)

16:L2a-M: Transformer mutual inductance ratio measurement for 11025

model use only. (L2A parameter is being set.)

17:L2b-M: Transformer mutual inductance ratio measurement for 11025

model use only. (L2B parameter is being set.)

18:L2a-R2:Transformer DC resistance measurement for 11025 model use

only. (L2A parameter is being set.)

19:L2b-R2:Transformer DC resistance measurement for 11025 model use

only. (L2B parameter is being set.)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetKeyLockState

ViStatus Chr1102x_SetKeyLockState (ViSession instrumentHandle,

ViBoolean keyLockState);

Purpose

It sets whether key lock or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

keyLockState

Variable Type       ViBoolean

It sets whether key lock or not.

0:Off (Unlock)

1:On  (Key Lock)

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetL2A_BParameter

ViStatus Chr1102x_SetL2A_BParameter (ViSession instrumentHandle,

ViBoolean l2A_BParameter);

Purpose

It sets the parameters of L2A and L2B.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

l2A_BParameter

Variable Type       ViBoolean

It sets the parameters of L2A and L2B.

Ture(1):L2A

False(0):L2B

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetLCDDisplay

ViStatus Chr1102x_SetLCDDisplay (ViSession instrumentHandle,

ViBoolean LCDDisplay);

Purpose

It sets LCD display is turned on or turn off status.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LCDDisplay

Variable Type       ViBoolean

It sets LCD display is turned on or turn off status.

Ture(1):ON

False(0):OFF

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetLCDText1Page

ViStatus Chr1102x_SetLCDText1Page (ViSession instrumentHandle,

ViInt32 LCDText1Page);

Purpose

It sets display screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LCDText1Page

Variable Type       ViInt32

It sets display screen.

Description:

1: Measuring screen

2: Comparison screen

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetLCDText2Page

ViStatus Chr1102x_SetLCDText2Page (ViSession instrumentHandle,

ViInt32 LCDText2Page);

Purpose

It sets page of measuring screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LCDText2Page

Variable Type       ViInt32

It sets page of measuring screen.

Description:

1: Show the first page of measuring screen (test frequency, level and

parameter).

2: Show the second page of measuring screen (file, average time and

trigger mode).

3: Show the third page of measuring screen (DC bias, binning and

comparison setting).

4:Show the four page of measuring screen (bias I state,Bias parameter

setting).

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetLengthofTestCable

ViStatus Chr1102x_SetLengthofTestCable (ViSession instrumentHandle,

ViInt32 lengthOfTestCable);

Purpose

It selects length of test cable.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

lengthOfTestCable

Variable Type       ViInt32

This parameter selects length of test cable.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetLINK1320Enable

ViStatus Chr1102x_SetLINK1320Enable (ViSession instrumentHandle,

ViBoolean LINK1320Enable);

Purpose

It sets LINK1320 ON/OFF.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

LINK1320Enable

Variable Type       ViBoolean

It sets LINK1320 ON/OFF.

1:On

0:Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetLowLimitofBin

ViStatus Chr1102x_SetLowLimitofBin (ViSession instrumentHandle,

ViInt32 binNumber,

ViReal64 lowLimitOfBIN);

Purpose

It sets the lower limit value of each BINs.  When BINning MODE is ABS, it

is the absolute value of inputted lower limit.  When BINning MODE is

PCNT, it is the percentage of inputted lower limit.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binNumber

Variable Type       ViInt32

It sets the number of BINs.

Valid Range: 1 ~ 8.

lowLimitOfBIN

Variable Type       ViReal64

It sets the lower limit value of each BINs.  When BINning MODE is

ABS, it is the absolute value of inputted lower limit.  When BINning

MODE is PCNT, it is the percentage of inputted lower limit.

Vaild Range:

ABS MODE: -999900000000000 ~ 999900000000000

PCNT MODE: -999.99 ~ 999.99

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetMeasurementSpeed

ViStatus Chr1102x_SetMeasurementSpeed (ViSession instrumentHandle,

ViReal64 measurementSpeed);

Purpose

It sets the measurement speed.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

measurementSpeed

Variable Type       ViReal64

It sets the measurement speed.

Fast:0.025

Medium:0.065

Slow:0.5

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetMeasureParallelMode

ViStatus Chr1102x_SetMeasureParallelMode (ViSession instrumentHandle,

ViInt32 setMeasureParallelMode);

Purpose

It sets circuit measurement parallel mode.

Note:Only for 11025

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setMeasureParallelMode

Variable Type       ViInt32

It sets circuit measurement parallel mode.

0:"FIMPedance"    Impedance measurement (equivalent series

circuit),"Set Measure Transformer State" will be set "off"

1:"FADMittance"   Admittance measurement (equivalent parallel

circuit),"Set Measure Transformer State" will be set "off"

2:'FIMPedance','FRESistance'  DCR measurement (equivalent series

circuit),"Set Measure Transformer State" will be set "ON"

3:'FADMittance','FRESistance' DCR measurement (equivalent parallel

circuit),"Set Measure Transformer State" will be set "ON"

4:'VOLTage:AC','IMPedance'    Transformer turn ratio measurement

,"Set Measure Transformer State" will be set "ON"

5:'IMPedance','FADMittance'   Transformer mutual inductance ratio

measurement,"Set Measure Transformer State" will be set "ON"

6:'IMPedance','RESistance'    Transformer DC resistance

measurement,"Set Measure Transformer State" will be set "ON"

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetMeasureRange

ViStatus Chr1102x_SetMeasureRange (ViSession instrumentHandle,

ViInt32 setMeasureRange);

Purpose

It sets the measurement file.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setMeasureRange

Variable Type       ViInt32

It sets the measurement file.

0:0.1 ohm

1:1 ohm

2:10 ohm

3:100 ohm

4:300 ohm  (only for 11022)

5:1k ohm

6:3k ohm  (only for 11022)

7:10k ohm

8:30k ohm  (only for 11022)

9:100k ohm

10:300k ohm  (only for 11022)

11:1M ohm

12:UP key simulation

13:Down key simulation

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetMeasureSpeedofDual

ViStatus Chr1102x_SetMeasureSpeedofDual (ViSession instrumentHandle,

ViReal64 setMeasureSpeedOfDual);

Purpose

It sets measure speed of dual mode.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setMeasureSpeedOfDual

Variable Type       ViReal64

It sets measure speed of dual mode.

Fast:0.025

Medium:0.065

Slow:0.5

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetMeasureTransformerState

ViStatus Chr1102x_SetMeasureTransformerState (ViSession instrumentHandle,

ViBoolean setMeasureTransformerState);

Purpose

It sets whether to measure transformer parameter (DCR, L2, N, M and R2)

or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setMeasureTransformerState

Variable Type       ViBoolean

It sets whether to measure transformer parameter (DCR, L2, N, M and

R2) or not.

Ture(1):ON

False(0):Off

ON allows two commands of <SENSOR_FUNCTION>, it means to

perform transformer measurement.

OFF allows one command of <SENSOR_FUNCTION>, it means to

perform common measurement.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetNominalofBin

ViStatus Chr1102x_SetNominalofBin (ViSession instrumentHandle,

ViReal64 setNominal);

Purpose

It sets NOMINAL value of BINNING, that is the nominal value of

measurement primary parameter A.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setNominal

Variable Type       ViReal64

It sets NOMINAL value of BINNING, that is the nominal value of

measurement parameter A.

Vaild Range: 0 ~ 999900000000000.0

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetNominalValueofCompare

ViStatus Chr1102x_SetNominalValueofCompare (ViSession instrumentHandle,

ViBoolean selectParameter,

ViReal64 nominalValueOfCompare);

Purpose

It sets primary, secondary parameter the nominal value in COMPARE mode.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

selectParameter

Variable Type       ViBoolean

Which parameter will be setting.

Set current (Im):0

Set voltage (Vm):1

nominalValueOfCompare

Variable Type       ViReal64

This parameter returns primary, secondary parameter the nominal value

in COMPARE mode.

Range:-999900000000000~999900000000000

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetNumericalRepresentMethod

ViStatus Chr1102x_SetNumericalRepresentMethod

(ViSession instrumentHandle, ViBoolean calculateNumber,

ViBoolean numericalRepresentMethod);

Purpose

It sets numerical represent method.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

numericalRepresentMethod

Variable Type       ViBoolean

It sets or queries numerical represent method.

DEV (0) :Bias absolute value.

PCNT (1): Percentage (PCNT).

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetNumRepresentMethodAction

ViStatus Chr1102x_SetNumRepresentMethodAction

(ViSession instrumentHandle, ViBoolean calculateNumber,

ViBoolean numRepresentMethodAction);

Purpose

It sets definition operation management is action or not.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

numRepresentMethodAction

Variable Type       ViBoolean

It sets definition operation management is action or not.

Description:

ON (1)    :  Operation management is action.

OFF (0)   :  Operation management is not action.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetOPC

ViStatus Chr1102x_SetOPC (ViSession instrumentHandle);

Purpose

When completed all operation, notify the 11022/11025 set ASCII character

1 to output queue.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetOuputImpedance

ViStatus Chr1102x_SetOuputImpedance (ViSession instrumentHandle,

ViInt32 ouputImpedance);

Purpose

It sets output impedance mode.  There are 100/25 ohm, 100 ohm, 10 ohm/CC,

25 ohm etc.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

ouputImpedance

Variable Type       ViInt32

It sets output impedance mode.  There are 100/25 ohm, 100 ohm, 10

ohm/CC, 25 ohm etc.

0:100/25 ohm

1:100 ohm

2:10 ohm/CC

3:25 ohm

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetPara1ofMeasurement

ViStatus Chr1102x_SetPara1ofMeasurement (ViSession instrumentHandle,

ViInt32 para1OfMeasurement);

Purpose

It sets the current measurement master parameter.

Description: REAL           Impedance real

MLINear        Impedance absolute value

CP             Equivalent parallel capacity

CS             Equivalent series capacity

LP             Equivalent parallel induction

LS             Equivalent series induction

ZS             Equivalent series impedance

RS             Equivalent series resistance

RP             Equivalent parallel resistance

NONE           DCR Measurement only.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

para1OfMeasurement

Variable Type       ViInt32

It sets the current measurement master parameter.

Description:

REAL           Impedance real                  :0

MLINear        Impedance absolute value        :1

CP             Equivalent parallel capacity    :2

CS             Equivalent series capacity      :3

LP             Equivalent parallel induction   :4

LS             Equivalent series induction     :5

ZS             Equivalent series impedance     :6

RS             Equivalent series resistance    :7

RP             Equivalent parallel resistance  :8

NONE           DCR Measurement only.           :9

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetPara2ofMeasurement

ViStatus Chr1102x_SetPara2ofMeasurement (ViSession instrumentHandle,

ViInt32 para2OfMeasurement);

Purpose

It sets the current measurement secondary parameter.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

para2OfMeasurement

Variable Type       ViInt32

It sets the current measurement secondary parameter.

Description:

IMAGinary          Impedance imaginary                 :0

PHASe         Phase                               :1

D             Dissipation factor                  :2

Q             Quality factor                      :3

REAL          Impedance real                      :4

RS            Equivalent parallel resistance      :5

XS            Imaginary series equivalent value   :6

INV           Inversion of transformer turn ratio :7

LP            Mutual     inductance               :8

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetPRI_HI_LIMITofCompare

ViStatus Chr1102x_SetPRI_HI_LIMITofCompare (ViSession instrumentHandle,

ViBoolean calculateNumber,

ViReal64 setPRI_HI_LIMITOfCompare);

Purpose

It sets the high limit value.

Description:

The low limit value needs between -9.999E14 ~ 9.999E14.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

setPRI_HI_LIMITOfCompare

Variable Type       ViReal64

It sets the high limit value.

Description:

The high limit value needs between -9.999E14 ~ 9.999E14.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetPRI_LO_LIMITofCompare

ViStatus Chr1102x_SetPRI_LO_LIMITofCompare (ViSession instrumentHandle,

ViBoolean calculateNumber,

ViReal64 setPRI_LO_LIMITOfCompare);

Purpose

It sets the low limit value.

Description:

The low limit value needs between -9.999E14 ~ 9.999E14.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Calculate 1: The master parameter.

Calculate 2: The secondary parameter.

setPRI_LO_LIMITOfCompare

Variable Type       ViReal64

This parameter sets the low limit value.

Description:

The low limit value needs between -9.999E14 ~ 9.999E14.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetRecallNumber

ViStatus Chr1102x_SetRecallNumber (ViSession instrumentHandle,

ViInt32 recallNumber);

Purpose

It recall data from instrument.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

recallNumber

Variable Type       ViInt32

It recall data from instrument.

Valid range is 0~ 49.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetRST

ViStatus Chr1102x_SetRST (ViSession instrumentHandle);

Purpose

It sets the 11022/11025 return preset status

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetSaveNumber

ViStatus Chr1102x_SetSaveNumber (ViSession instrumentHandle,

ViInt32 saveNumber);

Purpose

It save data to instrument.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

saveNumber

Variable Type       ViInt32

It save data to instrument.

Valid range is 0~ 49.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetSRE

ViStatus Chr1102x_SetSRE (ViSession instrumentHandle, ViInt32 setSRE);

Purpose

It sets start bit in register.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setSRE

Variable Type       ViInt32

It sets start bit in register.

Valid Range:0 ~ 255.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetSystemIntegration

ViStatus Chr1102x_SetSystemIntegration (ViSession instrumentHandle,

ViInt32 setSystemIntegration);

Purpose

It calculates period selection function for deciding every sampling

period when quick measuring.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setSystemIntegration

Variable Type       ViInt32

It calculates period selection function for deciding every sampling

period when quick measuring.

Range is 1~8.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTestFrequency

ViStatus Chr1102x_SetTestFrequency (ViSession instrumentHandle,

ViInt32 testFrequency);

Purpose

It sets test frequency now.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

testFrequency

Variable Type       ViInt32

It sets the test frequency now.

50:50 Hz

60:60 Hz

100:100 Hz

120:120 Hz

1000:1K Hz

10000:10K Hz

20000:20K Hz

40000:40K Hz

50000:50K Hz

100000:100K Hz

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTestVoltage

ViStatus Chr1102x_SetTestVoltage (ViSession instrumentHandle,

ViReal64 testVoltage);

Purpose

It sets test voltage now.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

testVoltage

Variable Type       ViReal64

It sets test voltage now.

Value Range :0.01~1 V

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTrigger

ViStatus Chr1102x_SetTrigger (ViSession instrumentHandle);

Purpose

It triggers the 1102x when the trigger mode is set to bus trigger mode.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTriggerDelayTime

ViStatus Chr1102x_SetTriggerDelayTime (ViSession instrumentHandle,

ViReal64 triggerDelayTime);

Purpose

It sets trigger delay time.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

triggerDelayTime

Variable Type       ViReal64

It sets trigger delay time.

The range is 0 ~ 9.999 s.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTriggerEdge

ViStatus Chr1102x_SetTriggerEdge (ViSession instrumentHandle,

ViInt32 triggerEdge);

Purpose

It sets TRIGGER mode to EXTERNAL, is RISING or FALLING trigger.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

triggerEdge

Variable Type       ViInt32

It sets TRIGGER mode to EXTERNAL, is RISING or FALLING trigger.

False(0):Falling

Ture(1):Rising

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTriggerSource

ViStatus Chr1102x_SetTriggerSource (ViSession instrumentHandle,

ViInt32 triggerSource);

Purpose

It sets trigger mode now.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

triggerSource

Variable Type       ViInt32

It sets or queries trigger mode now.

0:BUS Trigger

1:External Trigger

2:Internal Trigger

3:Manual Trigger

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTriggerSourceofDual

ViStatus Chr1102x_SetTriggerSourceofDual (ViSession instrumentHandle,

ViInt32 setTriggerSourceOfDual);

Purpose

It sets  trigger mode of dual now.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setTriggerSourceOfDual

Variable Type       ViInt32

It sets  trigger mode of dual now.

Description:

0:BUS Trigger

1:External Trigger

2:Internal Trigger

3:Manual Trigger

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetTriggerState

ViStatus Chr1102x_SetTriggerState (ViSession instrumentHandle);

Purpose

Whatever the measurement status now, trigger execution measurement

Warning1:

When users give TRIGger command, 11022/11025 measuring is activated.

Meanwhile,if FETCh? command is given, for taking a period of the actual

measurement time is needed then the measurement result fetched by FETCh?

Command will be shown. The calculation of this period of measurement time

is as the following:

Measurement Time= Trigger Delay+Average Time51ms= 202ms

Warning2:

When users give TRIGger command, the sequential other setting commands

(non FETCh ? commands) will be given. The program needs to write the

function of time delay. This delay time is the measurement time of

11022/11025, otherwise the part of

sequential commands cant normal function. Take above item for example,

it is needed to write 202ms delay time program to ensure the sequential

commands are all normal functioning.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetUpperLimitofBin

ViStatus Chr1102x_SetUpperLimitofBin (ViSession instrumentHandle,

ViInt32 binNumber,

ViReal64 upperLimitOfBIN);

Purpose

It sets the upper limit value of each BINs.  When BINning MODE is ABS, it

is the absolute value of inputted upper limit.  When BINning MODE is

PCNT, it is the percentage of inputted upper limit.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

binNumber

Variable Type       ViInt32

It sets the number of BINs.

Valid Range: 1 ~ 8.

upperLimitOfBIN

Variable Type       ViReal64

It sets the upper limit value of each BINs.  When BINning MODE is

ABS, it is the absolute value of inputted upper limit.  When BINning

MODE is PCNT, it is the percentage of inputted upper limit.

Vaild Range:

ABS MODE: -999900000000000 ~ 999900000000000

PCNT MODE: -999.99 ~ 999.99

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetVm_ImDisplayState

ViStatus Chr1102x_SetVm_ImDisplayState (ViSession instrumentHandle,

ViBoolean calculateNumber,

ViBoolean VM_IMDisplayState);

Purpose

It sets whether turn on Vm/Im function.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

calculateNumber

Variable Type       ViBoolean

Which parameter will be setting.

Set current (Im):0

Set voltage (Vm):1

VM_IMDisplayState

Variable Type       ViBoolean

It sets whether turn on Vm/Im function.

Description:

ON (1)        Turn on

OFF (0)       Turn off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_SetVoltageOffsetstate

ViStatus Chr1102x_SetVoltageOffsetstate (ViSession instrumentHandle,

ViBoolean setVoltageOffsetState);

Purpose

It sets if BIAS V function is enabled.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

setVoltageOffsetState

Variable Type       ViBoolean

It sets if BIAS V function is enabled.

Ture(1):On

False(0):Off

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_TriggerSourceofDual

ViStatus Chr1102x_TriggerSourceofDual (ViSession instrumentHandle,

ViChar _VI_FAR triggerSourceOfDual[]);

Purpose

It queries  trigger mode of dual now.

NOTE:

DUAL command set is only for dual frequency function screen that is

entering this screen.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

triggerSourceOfDual

Variable Type       ViChar[]

This parameter returns master parameter of first measure of the dual

function.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFC0103  Self-Test not supported.

BFFC0012  Error interpreting instrument response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_UnlockSession

ViStatus Chr1102x_UnlockSession (ViSession instrumentHandle,

ViPBoolean callerHasLock);

Purpose

This function releases a lock that you acquired on an instrument session

using Chr1102x_LockSession.  Refer to Chr1102x_LockSession for additional

information on session locks.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

callerHasLock

Variable Type       ViBoolean (passed by reference)

This parameter serves as a convenience.  If you do not want to use

this parameter, pass VI_NULL.

Use this parameter in complex functions to keep track of whether you

obtain a lock and therefore need to unlock the session.

Pass the address of a local ViBoolean variable.  In the declaration

of the local variable, initialize it to VI_FALSE.  Pass the address

of the same local variable to any other calls you make to

Chr1102x_LockSession or Chr1102x_UnlockSession in the same function.

The parameter is an input/output parameter.  Chr1102x_LockSession and

Chr1102x_UnlockSession each inspect the current value and take the

following actions:

- If the value is VI_TRUE, Chr1102x_LockSession does not lock the

session again.  If the value is VI_FALSE, Chr1102x_LockSession

obtains the lock and sets the value of the parameter to VI_TRUE.

- If the value is VI_FALSE, Chr1102x_UnlockSession does not attempt

to unlock the session.  If the value is VI_TRUE,

Chr1102x_UnlockSession releases the lock and sets the value of the

parameter to VI_FALSE.

Thus, you can, call Chr1102x_UnlockSession at the end of your

function without worrying about whether you actually have the lock.

Example:

ViStatus TestFunc (ViSession vi, ViInt32 flags)

{

ViStatus error = VI_SUCCESS;

ViBoolean haveLock = VI_FALSE;

if (flags & BIT_1)

{

viCheckErr( Chr1102x_LockSession(vi, &haveLock));

viCheckErr( TakeAction1(vi));

if (flags & BIT_2)

{

viCheckErr( Chr1102x_UnlockSession(vi, &haveLock));

viCheckErr( TakeAction2(vi));

viCheckErr( Chr1102x_LockSession(vi, &haveLock);

}

if (flags & BIT_3)

viCheckErr( TakeAction3(vi));

}

Error:

/*

At this point, you cannot really be sure that

you have the lock.  Fortunately, the haveLock

variable takes care of that for you.

*/

Chr1102x_UnlockSession(vi, &haveLock);

return error;

}

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Chr1102x_WriteInstrData

ViStatus Chr1102x_WriteInstrData (ViSession instrumentHandle,

ViChar _VI_FAR writeBuffer[]);

Purpose

This function writes a user-specified string to the instrument.

Note:  This function bypasses IVI attribute state caching.  Therefore,

when you call this function, the cached values for all attributes will be

invalidated.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the Chr1102x_init or

Chr1102x_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

writeBuffer

Variable Type       ViChar[]

Pass the string to be written to the instrument.

Return Value

Reports the status of this operation.

To obtain a text description of the status code, or if the status

code is not listed below, call the Chr1102x_error_message function.

To obtain additional information about the error condition, use the

Chr1102x_GetErrorInfo and Chr1102x_ClearErrorInfo functions.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  A "no listeners" condition was detected.

BFFF0060  This interface is not the controller-in-charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

Attribute Information for the Following Functions:

Chr1102x_SetAttributeViInt32

Chr1102x_GetAttributeViInt32

Chr1102x_CheckAttributeViInt32

Chr1102x_SetAttributeViReal64

Chr1102x_GetAttributeViReal64

Chr1102x_CheckAttributeViReal64

Chr1102x_SetAttributeViSession

Chr1102x_GetAttributeViSession

Chr1102x_CheckAttributeViSession

Chr1102x_SetAttributeViBoolean

Chr1102x_GetAttributeViBoolean

Chr1102x_CheckAttributeViBoolean

Chr1102x_SetAttributeViString

Chr1102x_GetAttributeViString

Chr1102x_CheckAttributeViString

Inherent IVI Attributes

User Options

Range Check                              CHR1102X_ATTR_RANGE_CHECK

Query Instrument Status                  CHR1102X_ATTR_QUERY_INSTR_STATUS

Cache                                    CHR1102X_ATTR_CACHE

Simulate                                 CHR1102X_ATTR_SIMULATE

Record Value Coercions                   CHR1102X_ATTR_RECORD_COERCIONS

Driver Identification

Description                              CHR1102X_ATTR_SPECIFIC_DRIVER_DESCRIPTION

Driver Prefix                            CHR1102X_ATTR_SPECIFIC_DRIVER_PREFIX

Driver Vendor                            CHR1102X_ATTR_SPECIFIC_DRIVER_VENDOR

Major Version                            CHR1102X_ATTR_SPECIFIC_DRIVER_MAJOR_VERSION

Minor Version                            CHR1102X_ATTR_SPECIFIC_DRIVER_MINOR_VERSION

Revision                                 CHR1102X_ATTR_SPECIFIC_DRIVER_REVISION

Class Specification Major Version        CHR1102X_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MAJOR_VERSION

Class Specification Minor Version        CHR1102X_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MINOR_VERSION

Driver Capabilities

Supported Instrument Models              CHR1102X_ATTR_SUPPORTED_INSTRUMENT_MODELS

Class Group Capabilities                 CHR1102X_ATTR_GROUP_CAPABILITIES

Number of Channels                       CHR1102X_ATTR_NUM_CHANNELS

Instrument Identification

Manufacturer                             CHR1102X_ATTR_INSTRUMENT_MANUFACTURER

Model                                    CHR1102X_ATTR_INSTRUMENT_MODEL

Firmware Revision                        CHR1102X_ATTR_INSTRUMENT_FIRMWARE_REVISION

Error Info

Primary Error                            CHR1102X_ATTR_PRIMARY_ERROR

Secondary Error                          CHR1102X_ATTR_SECONDARY_ERROR

Error Elaboration                        CHR1102X_ATTR_ERROR_ELABORATION

Advanced Session Information

Logical Name                             CHR1102X_ATTR_LOGICAL_NAME

Resource Descriptor                      CHR1102X_ATTR_RESOURCE_DESCRIPTOR

I/O Session Type                         CHR1102X_ATTR_IO_SESSION_TYPE

Instrument I/O Session                   CHR1102X_ATTR_IO_SESSION

Miscellaneous Attributes

CHR1102X_ATTR_MODEL                         CHR1102X_ATTR_MODEL

ID Query Response                           CHR1102X_ATTR_ID_QUERY_RESPONSE

Common

CHR1102X_ATTR_NAOMINAL                      CHR1102X_ATTR_NAOMINAL

CHR1102X_ATTR_SAVE_NUMBER                   CHR1102X_ATTR_SAVE_NUMBER

CHR1102X_ATTR_RECALL_NUMBER                 CHR1102X_ATTR_RECALL_NUMBER

CHR1102X_ATTR_STANDARD_EVENT_STATUS_ENABLE  CHR1102X_ATTR_STANDARD_EVENT_STATUS_ENABLE

CHR1102X_ATTR_OPC_EVENT                     CHR1102X_ATTR_OPC_EVENT

CHR1102X_ATTR_SRE_EVENT                     CHR1102X_ATTR_SRE_EVENT

CHR1102X_ATTR_BINNING_MODE                  CHR1102X_ATTR_BINNING_MODE

CHR1102X_ATTR_AUX_HI_OF_BIN                 CHR1102X_ATTR_AUX_HI_OF_BIN

CHR1102X_ATTR_LOW_LIMIT_OF_BIN              CHR1102X_ATTR_LOW_LIMIT_OF_BIN

CHR1102X_ATTR_AUX_LO_OF_BIN                 CHR1102X_ATTR_AUX_LO_OF_BIN

CHR1102X_ATTR_BINNING_GREATER_THAN_9_STATE  CHR1102X_ATTR_BINNING_GREATER_THAN_9_STATE

CHR1102X_ATTR_BINNING_NO_MORE_9_HI_VALUE    CHR1102X_ATTR_BINNING_NO_MORE_9_HI_VALUE

CHR1102X_ATTR_BINNING_NO_MORE_9_LO_VALUE    CHR1102X_ATTR_BINNING_NO_MORE_9_LO_VALUE

CHR1102X_ATTR_BINNING_NO_MORE_9             CHR1102X_ATTR_BINNING_NO_MORE_9

CHR1102X_ATTR_BIN_NUMBER                    CHR1102X_ATTR_BIN_NUMBER

CHR1102X_ATTR_BINNING_STATE                 CHR1102X_ATTR_BINNING_STATE

CHR1102X_ATTR_PARA_1_OF_MEASUREMENT         CHR1102X_ATTR_PARA_1_OF_MEASUREMENT

CHR1102X_ATTR_PARA_2_OF_MEASUREMENT         CHR1102X_ATTR_PARA_2_OF_MEASUREMENT

CHR1102X_ATTR_BUZZER_OUTPUT_DEFINE          CHR1102X_ATTR_BUZZER_OUTPUT_DEFINE

CHR1102X_ATTR_BUZZER_OUTPUT_ENABLE          CHR1102X_ATTR_BUZZER_OUTPUT_ENABLE

CHR1102X_ATTR_PRI_LO_LIMIT_OF_COMPARE       CHR1102X_ATTR_PRI_LO_LIMIT_OF_COMPARE

CHR1102X_ATTR_COMPARE_LIMIT_ENABLE          CHR1102X_ATTR_COMPARE_LIMIT_ENABLE

CHR1102X_ATTR_PRI_HI_LIMIT_OF_COMPARE       CHR1102X_ATTR_PRI_HI_LIMIT_OF_COMPARE

CHR1102X_ATTR_NUMERICAL_REPRESENT_METHOD    CHR1102X_ATTR_NUMERICAL_REPRESENT_METHOD

CHR1102X_ATTR_NUMERICAL_REPRESENT_METHOD_STATE  CHR1102X_ATTR_NUMERICAL_REPRESENT_METHOD_STATE

CHR1102X_ATTR_Vm_Im_MATH_STATE              CHR1102X_ATTR_Vm_Im_MATH_STATE

CHR1102X_ATTR_LENGEH_OF_TEST_CABLE          CHR1102X_ATTR_LENGEH_OF_TEST_CABLE

DATA

CHR1102X_ATTR_NOMINAL_VALUE_OF_COMPARE      CHR1102X_ATTR_NOMINAL_VALUE_OF_COMPARE

CHR1102X_ATTR_DATA_VALUE_OF_COMPARE         CHR1102X_ATTR_DATA_VALUE_OF_COMPARE

Display

CHR1102X_ATTR_LCD_DISPLAY_STATE             CHR1102X_ATTR_LCD_DISPLAY_STATE

CHR1102X_ATTR_LCD_TEXT1_PAGE                CHR1102X_ATTR_LCD_TEXT1_PAGE

CHR1102X_ATTR_LCD_TEXT2_PAGE                CHR1102X_ATTR_LCD_TEXT2_PAGE

DUAL

CHR1102X_ATTR_PARA_1_OF_DUAL_FIRST_MEASUREMENT  CHR1102X_ATTR_PARA_1_OF_DUAL_FIRST_MEASUREMENT

CHR1102X_ATTR_PARA_2_OF_DUAL_FIRST_MEASUREMENT  CHR1102X_ATTR_PARA_2_OF_DUAL_FIRST_MEASUREMENT

CHR1102X_ATTR_PARA_2_OF_DUAL_SECONDARY_MEASUREMENT  CHR1102X_ATTR_PARA_2_OF_DUAL_SECONDARY_MEASUREMENT

CHR1102X_ATTR_AUTO_STATE_OF_DUAL_FIRST_MEASUREMENT  CHR1102X_ATTR_AUTO_STATE_OF_DUAL_FIRST_MEASUREMENT

CHR1102X_ATTR_DUAL_FIRST_MEASURE_RANGE      CHR1102X_ATTR_DUAL_FIRST_MEASURE_RANGE

CHR1102X_ATTR_AUTO_STATE_OF_DUAL_SECOND_MEASUREMENT  CHR1102X_ATTR_AUTO_STATE_OF_DUAL_SECOND_MEASUREMENT

CHR1102X_ATTR_DUAL_SECOND_MEASURE_RANGE     CHR1102X_ATTR_DUAL_SECOND_MEASURE_RANGE

CHR1102X_ATTR_DUAL_FIRST_SOURCE_FREQUENCY   CHR1102X_ATTR_DUAL_FIRST_SOURCE_FREQUENCY

CHR1102X_ATTR_DUAL_SECOND_SOURCE_FREQUENCY  CHR1102X_ATTR_DUAL_SECOND_SOURCE_FREQUENCY

CHR1102X_ATTR_DUAL_BINNING_MODE             CHR1102X_ATTR_DUAL_BINNING_MODE

CHR1102X_ATTR_DUAL_FIRST_BINNING_NOMINAL    CHR1102X_ATTR_DUAL_FIRST_BINNING_NOMINAL

CHR1102X_ATTR_DUAL_FIRST_BINNING_AUX_HI     CHR1102X_ATTR_DUAL_FIRST_BINNING_AUX_HI

CHR1102X_ATTR_DUAL_FIRST_BINNING_AUX_LO     CHR1102X_ATTR_DUAL_FIRST_BINNING_AUX_LO

CHR1102X_ATTR_DUAL_SECOND_BINNING_AUX_HI    CHR1102X_ATTR_DUAL_SECOND_BINNING_AUX_HI

CHR1102X_ATTR_DUAL_SECOND_BINNING_AUX_LO    CHR1102X_ATTR_DUAL_SECOND_BINNING_AUX_LO

CHR1102X_ATTR_DUAL_FIRST_BIN_NUMBER         CHR1102X_ATTR_DUAL_FIRST_BIN_NUMBER

CHR1102X_ATTR_DUAL_FIRST_BIN_UPPER          CHR1102X_ATTR_DUAL_FIRST_BIN_UPPER

CHR1102X_ATTR_DUAL_FIRST_BIN_LOWER          CHR1102X_ATTR_DUAL_FIRST_BIN_LOWER

CHR1102X_ATTR_TRIGGER_SOURCE_OF_DUAL        CHR1102X_ATTR_TRIGGER_SOURCE_OF_DUAL

CHR1102X_ATTR_DUAL_MEASURE_SPEED            CHR1102X_ATTR_DUAL_MEASURE_SPEED

CHR1102X_ATTR_CONTINUOUS_INITIATE           CHR1102X_ATTR_CONTINUOUS_INITIATE

SENSE

CHR1102X_ATTR_AVERAGE_COUNT                 CHR1102X_ATTR_AVERAGE_COUNT

CHR1102X_ATTR_OPEN_SHORT_CORRECTION_PROCEDURE  CHR1102X_ATTR_OPEN_SHORT_CORRECTION_PROCEDURE

CHR1102X_ATTR_QUERY_CORRECTION_PROCEDURE    CHR1102X_ATTR_QUERY_CORRECTION_PROCEDURE

CHR1102X_ATTR_OPEN_CORRECTION_VALUE_G       CHR1102X_ATTR_OPEN_CORRECTION_VALUE_G

CHR1102X_ATTR_OPEN_CORRECTION_VALUE_B       CHR1102X_ATTR_OPEN_CORRECTION_VALUE_B

CHR1102X_ATTR_SHORT_CORRECTION_VALUE_R      CHR1102X_ATTR_SHORT_CORRECTION_VALUE_R

CHR1102X_ATTR_SHORT_CORRECTION_VALUE_X      CHR1102X_ATTR_SHORT_CORRECTION_VALUE_X

CHR1102X_ATTR_SHORT_1_OHM_CORRECTION_VALUE_R  CHR1102X_ATTR_SHORT_1_OHM_CORRECTION_VALUE_R

CHR1102X_ATTR_SHORT_1_OHM_CORRECTION_VALUE_X  CHR1102X_ATTR_SHORT_1_OHM_CORRECTION_VALUE_X

CHR1102X_ATTR_SHORT_100M_OHM_CORRECTION_VALUE_R  CHR1102X_ATTR_SHORT_100M_OHM_CORRECTION_VALUE_R

CHR1102X_ATTR_SHORT_100M_OHM_CORRECTION_VALUE_X  CHR1102X_ATTR_SHORT_100M_OHM_CORRECTION_VALUE_X

CHR1102X_ATTR_MEASURE_SPEED                 CHR1102X_ATTR_MEASURE_SPEED

CHR1102X_ATTR_AUTO_RANGE_STATE              CHR1102X_ATTR_AUTO_RANGE_STATE

CHR1102X_ATTR_MEASURE_RANGE                 CHR1102X_ATTR_MEASURE_RANGE

CHR1102X_ATTR_MEASUREMENT_TRANSFORMAER_PARAMETER  CHR1102X_ATTR_MEASUREMENT_TRANSFORMAER_PARAMETER

CHR1102X_ATTR_MEASUREMENT_L2A_OR_L2B        CHR1102X_ATTR_MEASUREMENT_L2A_OR_L2B

CHR1102X_ATTR_MEASURE_PARALLEL_MODE         CHR1102X_ATTR_MEASURE_PARALLEL_MODE

SOURCE

CHR1102X_ATTR_TEST_FREQUENCY                CHR1102X_ATTR_TEST_FREQUENCY

CHR1102X_ATTR_TEST_VOLTAGE                  CHR1102X_ATTR_TEST_VOLTAGE

CHR1102X_ATTR_BIAS_CURRENT_STATE            CHR1102X_ATTR_BIAS_CURRENT_STATE

CHR1102X_ATTR_VOLTAGE_OFFSET_STATE          CHR1102X_ATTR_VOLTAGE_OFFSET_STATE

CHR1102X_ATTR_BIAS_CURRENT_VALUE            CHR1102X_ATTR_BIAS_CURRENT_VALUE

SYSTEM

CHR1102X_ATTR_BEEP_STATE                    CHR1102X_ATTR_BEEP_STATE

CHR1102X_ATTR_KEY_LOCK_STATE                CHR1102X_ATTR_KEY_LOCK_STATE

CHR1102X_ATTR_OUTPUT_IMPEDANCE              CHR1102X_ATTR_OUTPUT_IMPEDANCE

CHR1102X_ATTR_SYSTEM_INTEGRATION            CHR1102X_ATTR_SYSTEM_INTEGRATION

CHR1102X_ATTR_BUZZER_ACTION_TYPE            CHR1102X_ATTR_BUZZER_ACTION_TYPE

CHR1102X_ATTR_HANDLER_MODE                  CHR1102X_ATTR_HANDLER_MODE

CHR1102X_ATTR_LINK_1320_ENABLE              CHR1102X_ATTR_LINK_1320_ENABLE

TRIGGER

CHR1102X_ATTR_TRIGGER_DELAY_TIME            CHR1102X_ATTR_TRIGGER_DELAY_TIME

CHR1102X_ATTR_TRIGGER_STATE                 CHR1102X_ATTR_TRIGGER_STATE

CHR1102X_ATTR_TRIGGER_SOURCE_MODE           CHR1102X_ATTR_TRIGGER_SOURCE_MODE

CHR1102X_ATTR_TRIGGER_EDGE                  CHR1102X_ATTR_TRIGGER_EDGE

CHR1102X_ATTR_ENABLE_REGISTER_STATUS        CHR1102X_ATTR_ENABLE_REGISTER_STATUS

CHR1102X_ATTR_AUTO_RANGE_STATE

Data Type:     ViBoolean

Description:   It sets or queries whether turn on auto file selection mode.

CHR1102X_ATTR_AUTO_STATE_OF_DUAL_FIRST_MEASUREMENT

Data Type:     ViBoolean

Description:   It is auto state of first measure of the dual function.

1:ON

0:OFF

CHR1102X_ATTR_AUTO_STATE_OF_DUAL_SECOND_MEASUREMENT

Data Type:     ViBoolean

Description:   It is auto state of second measure of the dual function.

1:ON

0:OFF

CHR1102X_ATTR_AUX_HI_OF_BIN

Data Type:     ViReal64

Description:   It sets AUX-HI value of BINNING, that is the upper limit value

of measurement

parameter B.

CHR1102X_ATTR_AUX_LO_OF_BIN

Data Type:     ViReal64

Description:   It sets AUX-LO value of BINNING, that is the lower limit value

of

measurement parameter B.

CHR1102X_ATTR_AVERAGE_COUNT

Data Type:     ViInt32

Description:   It sets or queries average time.

CHR1102X_ATTR_BEEP_STATE

Data Type:     ViInt32

Description:   It sets or queries whether buzzer is turned on or not.

OFF (0) - turn off buzzer

ON (1) - turn on buzzer (small)

LARGE (2) - turn on buzzer (large)

CHR1102X_ATTR_BIAS_CURRENT_STATE

Data Type:     ViBoolean

Description:   It returns and sets the output condition of bias current source.

CHR1102X_ATTR_BIAS_CURRENT_VALUE

Data Type:     ViReal64

Description:   It returns and sets bias current value.

Description: Be able to set the range of bias current value of a 1320 Bias

Current Source

0-10A as 0.000A~10.000A.  Be able to set the range of bias current value

of a 1320 Bias Current Source 0-20A as 0.000A~20.000A.  When a 1320

Bias Current Source 0-20A connects a 1320S, bias current value range can

be set is 0.000A ~ 40.000A.  A 1320 Bias Current Source 0-20A can

connect four 1320S maximumly, thus bias current value can be set is

0.000A ~ 100.000A.

CHR1102X_ATTR_BIN_NUMBER

Data Type:     ViInt32

Description:   The bin number.

CHR1102X_ATTR_BINNING_GREATER_THAN_9_STATE

Data Type:     ViBoolean

Description:   It sets enable state when binning number greater than 9.

CHR1102X_ATTR_BINNING_MODE

Data Type:     ViBoolean

Description:   It sets MODE to absolute value or percentage.

CHR1102X_ATTR_BINNING_NO_MORE_9

Data Type:     ViInt32

Description:   It sets bin number of bin when binning number greater than 9.

CHR1102X_ATTR_BINNING_NO_MORE_9_HI_VALUE

Data Type:     ViReal64

Description:   It sets maximum value of bin when binning number greater than 9.

CHR1102X_ATTR_BINNING_NO_MORE_9_LO_VALUE

Data Type:     ViReal64

Description:   It sets minimum value of bin when binning number greater than 9.

CHR1102X_ATTR_BINNING_STATE

Data Type:     ViBoolean

Description:   It sets BINNING is open or not.

CHR1102X_ATTR_BUZZER_ACTION_TYPE

Data Type:     ViBoolean

Description:   It selects buzzer action is PULSE or CONTINUOUS when BINSORTING

and

COMPARE.

CHR1102X_ATTR_BUZZER_OUTPUT_DEFINE

Data Type:     ViBoolean

Description:   Buzzer output definition.

Description:

FAIL(0):When the comparison result is FAIL, buzzer will be action.

PASS(1):When the comparison result is PASS, buzzer will be action.

CHR1102X_ATTR_BUZZER_OUTPUT_ENABLE

Data Type:     ViBoolean

Description:   It sets or queries whether buzzer is action or not.

Description:

ON (1):    Turn on buzzer.

OFF (0):   Turn off buzzer.

CHR1102X_ATTR_CACHE

Data Type:     ViBoolean

Description:   Specifies whether to cache the value of attributes.  When

caching is enabled, the instrument driver keeps track of the current instrument

settings and avoids sending redundant commands to the instrument.  Thus, you

can significantly increase execution speed.

The instrument driver can choose always to cache or never to cache

particular attributes regardless of the setting of this attribute.

The default value is VI_TRUE.   Use the Chr1102x_InitWithOptions function

to override this value.

CHR1102X_ATTR_COMPARE_LIMIT_ENABLE

Data Type:     ViBoolean

Description:   It sets or queries whether the comparison function is action or

not.

Description:

ON(1)    Turn on comparison function

OFF(0)    Turn off comparison function

CHR1102X_ATTR_CONTINUOUS_INITIATE

Data Type:     ViBoolean

Description:   It is initial trigger system sequentially.

CHR1102X_ATTR_DATA_VALUE_OF_COMPARE

Data Type:     ViReal64

Description:   It queries REF1 (the nominal value of COMPARE in primary

parameter),

REF2 (the nominal value of COMPARE in secondary parameter),

IMON (measuring current value),

VMON (measuring voltage value),

When Im/Vm Display function is OFF, return values of IMON, VMON is

0.

CHR1102X_ATTR_DUAL_BINNING_MODE

Data Type:     ViBoolean

Description:   It is binning mode of the dual function.

0:ABS

1:PERCENT

CHR1102X_ATTR_DUAL_FIRST_BIN_LOWER

Data Type:     ViReal64

Description:   It is the upper limit value of each BINs of dual function.  When

BINning: MODE ABS, it is the absolute value of inputted upper limit.  When

BINning: MODE PCNT, it is the percentage of inputted upper limit.

Vaild Range: -99999999999999.0 ~ 99999999999999.0

CHR1102X_ATTR_DUAL_FIRST_BIN_NUMBER

Data Type:     ViInt32

Description:

CHR1102X_ATTR_DUAL_FIRST_BIN_UPPER

Data Type:     ViReal64

Description:   It is the upper limit value of each BINs of dual function.  When

BINning: MODE ABS, it is the absolute value of inputted upper limit.  When

BINning: MODE PCNT, it is the percentage of inputted upper limit.

Vaild Range: -99999999999999.0 ~ 99999999999999.0

CHR1102X_ATTR_DUAL_FIRST_BINNING_AUX_HI

Data Type:     ViReal64

Description:   It is first binning aux-Hi value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

CHR1102X_ATTR_DUAL_FIRST_BINNING_AUX_LO

Data Type:     ViReal64

Description:   It is first binning aux-Lo value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

CHR1102X_ATTR_DUAL_FIRST_BINNING_NOMINAL

Data Type:     ViReal64

Description:   It is first binning nominal value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

CHR1102X_ATTR_DUAL_FIRST_MEASURE_RANGE

Data Type:     ViString

Description:   It is range value of first measure of the dual function.

0:0.1 ohm

1:1 ohm

2:10 ohm

3:100 ohm

4:1k ohm

5:10k ohm

6:100k ohm

7:1M ohm

8:UP key simulation

9:Down key simulation

CHR1102X_ATTR_DUAL_FIRST_SOURCE_FREQUENCY

Data Type:     ViInt32

Description:   It is frequnecy range of first measure of the dual function.

50:50 Hz

60:60 Hz

100:100 Hz

120:120 Hz

1000:1K Hz

10000:10K Hz

20000:20K Hz

40000:40K Hz

50000:50K Hz

100000:100K Hz

CHR1102X_ATTR_DUAL_MEASURE_SPEED

Data Type:     ViReal64

Description:   It is measure speed of dual mode.

Fast:0.025

Medium:0.065

Slow:0.5

CHR1102X_ATTR_DUAL_SECOND_BINNING_AUX_HI

Data Type:     ViReal64

Description:   It is second binning aux-Hi value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

CHR1102X_ATTR_DUAL_SECOND_BINNING_AUX_LO

Data Type:     ViReal64

Description:   It is second binning aux-Lo value of the dual function.

Description:

The value needs between -9.999E14 ~ 9.999E14.

CHR1102X_ATTR_DUAL_SECOND_MEASURE_RANGE

Data Type:     ViString

Description:   It is range value of second measure of the dual function.

0:0.1 ohm

1:1 ohm

2:10 ohm

3:100 ohm

4:1k ohm

5:10k ohm

6:100k ohm

7:1M ohm

8:UP key simulation

9:Down key simulation

CHR1102X_ATTR_DUAL_SECOND_SOURCE_FREQUENCY

Data Type:     ViInt32

Description:   It is frequnecy range of second measure of the dual function.

50:50 Hz

60:60 Hz

100:100 Hz

120:120 Hz

1000:1K Hz

10000:10K Hz

20000:20K Hz

40000:40K Hz

50000:50K Hz

100000:100K Hz

CHR1102X_ATTR_ENABLE_REGISTER_STATUS

Data Type:     ViInt32

Description:   It sets enable register status.

Range:0~255

CHR1102X_ATTR_ERROR_ELABORATION

Data Type:     ViString

Description:   An optional string that contains additional information

concerning the primary error condition.

CHR1102X_ATTR_GROUP_CAPABILITIES

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains a comma-separated list of class-extention

groups that this driver implements.

CHR1102X_ATTR_HANDLER_MODE

Data Type:     ViBoolean

Description:   It selects HANDLER MODE.

CHR1102X_ATTR_ID_QUERY_RESPONSE

Data Type:     ViString

Restrictions:  Not settable.

Description:   Returns the ID Query response string.  The instrument driver

gets the value of this attribute when you pass VI_TRUE for the ID Query

parameter to the Chr1102x_init or Chr1102x_InitWithOptions function.

CHR1102X_ATTR_INSTRUMENT_FIRMWARE_REVISION

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the firmware revision information for the

instrument you are currently using.

CHR1102X_ATTR_INSTRUMENT_MANUFACTURER

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the name of the instrument manufacturer

you are currently using.

CHR1102X_ATTR_INSTRUMENT_MODEL

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the model number or name of the

instrument that you are currently using.

CHR1102X_ATTR_IO_SESSION

Data Type:     ViSession

Restrictions:  Not settable.

Description:   Specifies the I/O session that the specific driver uses to

communicate with the instrument.

CHR1102X_ATTR_IO_SESSION_TYPE

Data Type:     ViString

Restrictions:  Not settable.

Description:   Contains the type of the I/O session that this driver uses to

communicate with the instrument.

CHR1102X_ATTR_KEY_LOCK_STATE

Data Type:     ViBoolean

Description:   It sets or queries whether 1102x key lock or not.

CHR1102X_ATTR_LCD_DISPLAY_STATE

Data Type:     ViBoolean

Description:   It sets or queries LCD display is turned on or turn off status.

CHR1102X_ATTR_LCD_TEXT1_PAGE

Data Type:     ViInt32

Description:   It sets or queries display screen.

Description:

1: Measuring screen

2: Comparison screen

CHR1102X_ATTR_LCD_TEXT2_PAGE

Data Type:     ViInt32

Description:   It sets or queries page of measuring screen.

Description:

1. Show the first page of measuring screen (test frequency, level and

CHR1102X_ATTR_LENGEH_OF_TEST_CABLE

Data Type:     ViInt32

Description:   It selects length of test cable.

CHR1102X_ATTR_LINK_1320_ENABLE

Data Type:     ViBoolean

Description:   It sets LINK1320 ON/OFF.

CHR1102X_ATTR_LOGICAL_NAME

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string containing the logical name you specified when opening

the current IVI session.

You may pass a logical name to the Chr1102x_init or

Chr1102x_InitWithOptions functions.  The IVI Configuration utility must contain

an entry for the logical name.  The logical name entry refers to a virtual

instrument section in the IVI Configuration file.  The virtual instrument

section specifies a physical device and initial user options.

CHR1102X_ATTR_LOW_LIMIT_OF_BIN

Data Type:     ViReal64

Description:   It sets the lower limit value of each BINs.  When BINning: MODE

ABS, it is

CHR1102X_ATTR_MEASURE_PARALLEL_MODE

Data Type:     ViString

Description:   It sets or queries circuit measurement parallel mode.

CHR1102X_ATTR_MEASURE_RANGE

Data Type:     ViString

Description:   It sets or queries the measurement range.

CHR1102X_ATTR_MEASURE_SPEED

Data Type:     ViReal64

Description:   It sets or queries the measurement speed.

CHR1102X_ATTR_MEASUREMENT_L2A_OR_L2B

Data Type:     ViBoolean

Description:   It sets the measurement of L2 A or B parameter.

CHR1102X_ATTR_MEASUREMENT_TRANSFORMAER_PARAMETER

Data Type:     ViBoolean

Description:   It sets whether to measure transformer parameter (DCR, L2, N, M

and R2) or

CHR1102X_ATTR_MODEL

Data Type:     ViInt32

Description:

Values:

CHR1102X_VAL_MODEL_1102X_2                     0

CHR1102X_VAL_MODEL_1102X_5                     1

CHR1102X_ATTR_NAOMINAL

Data Type:     ViReal64

Description:

CHR1102X_ATTR_NOMINAL_VALUE_OF_COMPARE

Data Type:     ViReal64

Description:   It sets primary, secondary parameter the nominal value in

COMPARE mode.

CHR1102X_ATTR_NUM_CHANNELS

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   Indicates the number of channels that the specific instrument

driver supports.

For each attribute for which the IVI_VAL_MULTI_CHANNEL flag attribute is

set, the IVI engine maintains a separate cache value for each channel.

CHR1102X_ATTR_NUMERICAL_REPRESENT_METHOD

Data Type:     ViBoolean

Description:   It sets or queries numerical represent method

CHR1102X_ATTR_NUMERICAL_REPRESENT_METHOD_STATE

Data Type:     ViBoolean

Description:   It sets or queries method state is action or not.

Description: ON (1)      Operation management is action.

OFF (0)      Operation management is not action.

CHR1102X_ATTR_OPC_EVENT

Data Type:     ViInt32

Description:   When completed all operation, notify the 11025 Standard Event

Register to set bit 0.

CHR1102X_ATTR_OPEN_CORRECTION_VALUE_B

Data Type:     ViReal64

Description:   It is OPEN correction procedure B.

CHR1102X_ATTR_OPEN_CORRECTION_VALUE_G

Data Type:     ViReal64

Description:   It is OPEN correction procedure G.

CHR1102X_ATTR_OPEN_SHORT_CORRECTION_PROCEDURE

Data Type:     ViInt32

Description:   It performs OPEN, SHORT correction procedure.

CHR1102X_ATTR_OUTPUT_IMPEDANCE

Data Type:     ViString

Description:   It sets output impedance mode.  There are 100/25W, 100W, 10W/CC,

25W etc.

CHR1102X_ATTR_PARA_1_OF_DUAL_FIRST_MEASUREMENT

Data Type:     ViString

Description:   It is master parameter of first measure of the dual function.

0:REAL

1:MLINEAR

2:CP

3:CS

4:LP

5:LS

6:ZS

7:RS

8:RP

CHR1102X_ATTR_PARA_1_OF_MEASUREMENT

Data Type:     ViString

Description:   It sets or queries the current measurement master parameter.

Description: REAL           Impedance real

MLINear        Impedance absolute value

CP             Equivalent parallel capacity

CS             Equivalent series capacity

LP             Equivalent parallel induction

LS             Equivalent series induction

ZS             Equivalent series impedance

RS             Equivalent series resistance

RP             Equivalent parallel resistance

CHR1102X_ATTR_PARA_2_OF_DUAL_FIRST_MEASUREMENT

Data Type:     ViString

Description:   It sets secondary parameter of first measure of the dual

function.

0:Imaginary

1:Phase

2:D

3:Q

4:REAL

5:RS

6:XS

CHR1102X_ATTR_PARA_2_OF_DUAL_SECONDARY_MEASUREMENT

Data Type:     ViString

Description:   It sets secondary parameter of secondary measure of the dual

function.

0:Imaginary

1:Phase

2:D

3:Q

4:REAL

5:RS

6:XS

CHR1102X_ATTR_PARA_2_OF_MEASUREMENT

Data Type:     ViString

Description:   It sets or queries the current measurement secondary parameter.

Description: IMAGinary      Impedance imaginary

PHASe          Phase

D               Dissipation factor

Q               Quality factor

REAL           Impedance real

RS              Equivalent parallel resistance

XS              Imaginary series equivalent value

CHR1102X_ATTR_PRI_HI_LIMIT_OF_COMPARE

Data Type:     ViReal64

Description:   It sets or queries the high limit value.

Description:

The high limit value needs between -9.999E14 ~ 9.999E14

Down directly MAXimum (9.999E14) or MINimum (-9.999E14)

CHR1102X_ATTR_PRI_LO_LIMIT_OF_COMPARE

Data Type:     ViReal64

Description:   It sets or queries the low limit value.

Description: The low limit value needs between -9.999E14 ~ 9.999E14

Down directly MAXimum (9.999E14) or MINimum (-9.999E14)

CHR1102X_ATTR_PRIMARY_ERROR

Data Type:     ViInt32

Description:   A code that describes the first error that occurred since the

last call to Chr1102x_GetErrorInfo on the session.  The value follows the

VXIplug&play completion code conventions.  A negative value describes an error

condition.  A positive value describes a warning condition and indicates that

no error occurred.  A zero indicates that no error or warning occurred. The

error and warning values can be status codes defined by IVI, VISA, class

drivers, or specific drivers.

CHR1102X_ATTR_QUERY_CORRECTION_PROCEDURE

Data Type:     ViInt32

Description:   It queries correction numeral.

CHR1102X_ATTR_QUERY_INSTR_STATUS

Data Type:     ViBoolean

Description:   Specifies whether the instrument driver queries the instrument

status after each operation.  Querying the instrument status is very useful for

debugging.  After you validate your program, you can set this attribute to

VI_FALSE to disable status checking and maximize performance

The instrument driver can choose to ignore status checking for particular

attributes regardless of the setting of this attribute.

The default value is VI_TRUE.   Use the Chr1102x_InitWithOptions function

to override this value.

CHR1102X_ATTR_RANGE_CHECK

Data Type:     ViBoolean

Description:   Specifies whether to validate attribute values and function

parameters.  If enabled, the instrument driver validates the parameter values

that you pass to driver functions.  Range checking parameters is very useful

for debugging.  After you validate your program, you can set this attribute to

VI_FALSE to disable range checking and maximize performance.

The default value is VI_TRUE.   Use the Chr1102x_InitWithOptions function

to override this value.

CHR1102X_ATTR_RECALL_NUMBER

Data Type:     ViInt32

Description:   It recall data from instrument.Valid range is 0 ~49.

CHR1102X_ATTR_RECORD_COERCIONS

Data Type:     ViBoolean

Description:   Specifies whether the IVI engine keeps a list of the value

coercions it makes for ViInt32 and ViReal64 attributes.  You call

Chr1102x_GetNextCoercionRecord to extract and delete the oldest coercion record

from the list.

The default value is VI_FALSE.   Use the Chr1102x_InitWithOptions function

to override this value.

CHR1102X_ATTR_RESOURCE_DESCRIPTOR

Data Type:     ViString

Restrictions:  Not settable.

Description:   Indicates the resource descriptor the driver uses to identify

the physical device.

If you initialize the driver with a logical name, this attribute contains

the resource descriptor that corresponds to the entry in the IVI Configuration

utility.

If you initialize the instrument driver with the resource descriptor, this

attribute contains that value.

CHR1102X_ATTR_SAVE_NUMBER

Data Type:     ViInt32

Description:   It save data to instrument.Valid range is 0 ~49.

CHR1102X_ATTR_SECONDARY_ERROR

Data Type:     ViInt32

Description:   An optional code that provides additional information concerning

the primary error condition.  The error and warning values can be status codes

defined by IVI, VISA, class drivers, or specific drivers.  Zero indicates no

additional information.

CHR1102X_ATTR_SHORT_100M_OHM_CORRECTION_VALUE_R

Data Type:     ViReal64

Description:   It is SHORT correction procedure R when range is 100 mohm.

CHR1102X_ATTR_SHORT_100M_OHM_CORRECTION_VALUE_X

Data Type:     ViReal64

Description:   It is SHORT correction procedure X when range is 100 mohm.

CHR1102X_ATTR_SHORT_1_OHM_CORRECTION_VALUE_R

Data Type:     ViReal64

Description:   It is SHORT correction procedure R when range is 1 ohm.

CHR1102X_ATTR_SHORT_1_OHM_CORRECTION_VALUE_X

Data Type:     ViReal64

Description:   It is SHORT correction procedure X when range is 1 ohm.

CHR1102X_ATTR_SHORT_CORRECTION_VALUE_R

Data Type:     ViReal64

Description:   It is SHORT correction procedure R.

CHR1102X_ATTR_SHORT_CORRECTION_VALUE_X

Data Type:     ViReal64

Description:   It is SHORT correction procedure X.

CHR1102X_ATTR_SIMULATE

Data Type:     ViBoolean

Description:   Specifies whether or not to simulate instrument driver I/O

operations.  If simulation is enabled, instrument driver functions perform

range checking and call Ivi_GetAttribute and Ivi_SetAttribute functions, but

they do not perform instrument I/O.  For output parameters that represent

instrument data, the instrument driver functions return calculated values.

The default value is VI_FALSE.   Use the Chr1102x_InitWithOptions function

to override this value.

CHR1102X_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MAJOR_VERSION

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   The major version number of the class specification with which

this driver is compliant.

CHR1102X_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MINOR_VERSION

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   The minor version number of the class specification with which

this driver is compliant.

CHR1102X_ATTR_SPECIFIC_DRIVER_DESCRIPTION

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains a brief description of the specific

driver

CHR1102X_ATTR_SPECIFIC_DRIVER_MAJOR_VERSION

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   The major version number of this instrument driver.

CHR1102X_ATTR_SPECIFIC_DRIVER_MINOR_VERSION

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   The minor version number of this instrument driver.

CHR1102X_ATTR_SPECIFIC_DRIVER_PREFIX

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the prefix for the instrument driver.

The name of each user-callable function in this driver starts with this prefix.

CHR1102X_ATTR_SPECIFIC_DRIVER_REVISION

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains additional version information about this

instrument driver.

CHR1102X_ATTR_SPECIFIC_DRIVER_VENDOR

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the name of the vendor that supplies this

driver.

CHR1102X_ATTR_SRE_EVENT

Data Type:     ViInt32

Description:   It sets start bit in register.

CHR1102X_ATTR_STANDARD_EVENT_STATUS_ENABLE

Data Type:     ViInt32

Description:   It sets or queries the bits of the Standard Event Status Enable

Register.

Valid range is 0 ~ 255.

CHR1102X_ATTR_SUPPORTED_INSTRUMENT_MODELS

Data Type:     ViString

Restrictions:  Not settable.

Description:   Contains a model code of the instrument. For drivers that

support more than one device, this attribute contains a comma-separated list of

supported instrument models.

CHR1102X_ATTR_SYSTEM_INTEGRATION

Data Type:     ViInt32

Description:   It calculates period selection function for deciding every

sampling period when

quick measuring.  Range is 1~8.

CHR1102X_ATTR_TEST_FREQUENCY

Data Type:     ViInt32

Description:   It sets or queries the test frequency now.

CHR1102X_ATTR_TEST_VOLTAGE

Data Type:     ViReal64

Description:   It sets or queries the test voltage now.

CHR1102X_ATTR_TRIGGER_DELAY_TIME

Data Type:     ViReal64

Description:   It sets or queries trigger delay time, the range is 0 ~ 9999ms.

CHR1102X_ATTR_TRIGGER_EDGE

Data Type:     ViString

Description:   It sets TRIGGER mode to EXTERNAL, is RISING or FALLING trigger.

CHR1102X_ATTR_TRIGGER_SOURCE_MODE

Data Type:     ViString

Description:   It sets or queries trigger mode now.

Description:

BUS         BUS Trigger

EXTernal     External Trigger

INTernal      Internal Trigger

MANual      Manual Trigger

CHR1102X_ATTR_TRIGGER_SOURCE_OF_DUAL

Data Type:     ViString

Description:   It is  trigger mode of dual now.

Description:

0:BUS Trigger

1:External Trigger

2:Internal Trigger

3:Manual Trigger

CHR1102X_ATTR_TRIGGER_STATE

Data Type:     ViBoolean

Description:   Whatever the measurement status now, trigger execution

measurement.

CHR1102X_ATTR_Vm_Im_MATH_STATE

Data Type:     ViBoolean

Description:   It sets or queries whether turn on Vm/Im function.

CHR1102X_ATTR_VOLTAGE_OFFSET_STATE

Data Type:     ViBoolean

Description: