Chroma 17040E

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Chroma 17040E

library.  The functions appear in alphabetical order, with a description of

the function and its C syntax, a description of each parameter, and a list of

possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                                             Function Name:

Initialize                                                 chr17040E_init

Initialize With Options                                    chr17040E_InitWithOptions

Configuration Functions

Charge/Discharge Controller

Configure Source Type                                chr17040E_ConfigureSourceType

Configure CC Charge Current Level                    chr17040E_ConfigureCCChargeCurrentLevel

Configure CC Charge Voltage Condition                chr17040E_ConfigureCCChargeVoltageCondition

Configure CC Charge Power Limit                      chr17040E_ConfigureCCChargePowerLimit

Configure CV Charge Voltage Level                    chr17040E_ConfigureCVChargeVoltageLevel

Configure CV Charge Current Condition                chr17040E_ConfigureCVChargeCurrentCondition

Configure CV Charge Power Limit                      chr17040E_ConfigureCVChargePowerLimit

Configure CP Charge Power Level                      chr17040E_ConfigureCPChargePowerLevel

Configure CP Charge Voltage Condition                chr17040E_ConfigureCPChargeVoltageCondition

Configure CP Charge Current Condition                chr17040E_ConfigureCPChargeCurrentCondition

Configure CC Discharge Current Level                 chr17040E_ConfigureCCDischargeCurrentLevel

Configure CC Discharge Voltage Condition             chr17040E_ConfigureCCDischargeVoltageCondition

Configure CC Discharge Power Limit                   chr17040E_ConfigureCCDischargePowerLimit

Configure CV Discharge Voltage Level                 chr17040E_ConfigureCVDischargeVoltageLevel

Configure CV Discharge Current Condition             chr17040E_ConfigureCVDischargeCurrentCondition

Configure CV Discharge Power Limit                   chr17040E_ConfigureCVDischargePowerLimit

Configure CP Discharge Power Level                   chr17040E_ConfigureCPDischargePowerLevel

Configure CP Discharge Voltage Condition             chr17040E_ConfigureCPDischargeVoltageCondition

Configure CP Discharge Current Condition             chr17040E_ConfigureCPDischargeCurrentCondition

Configure CV Source Voltage Level                    chr17040E_ConfigureCVSourceVoltageLevel

Configure CV Source Current Limit                    chr17040E_ConfigureCVSourceCurrentLimit

Configure Cutoff Time                                chr17040E_ConfigureCutoffTime

Configure Current Slew Rate                          chr17040E_ConfigureCurrentSlewRate

Protection

Configure Over Voltage Protection                 chr17040E_ConfigureOverVoltageProtection

Configure Under Voltage Protection                chr17040E_ConfigureUnderVoltageProtection

Configure Positive Delta Voltage Protection       chr17040E_ConfigurePositiveDeltaVoltageProtection

Configure Negative Delta Voltage Protection       chr17040E_ConfigureNegativeDeltaVoltageProtection

Configure Over Current Protection                 chr17040E_ConfigureOverCurrentProtection

Configure Positive Delta Current Protection       chr17040E_ConfigurePositiveDeltaCurrentProtection

Configure Negative Delta Current Protection       chr17040E_ConfigureNegativeDeltaCurrentProtection

Configure Over Capacity Protection                chr17040E_ConfigureOverCapacityProtection

Specification Query

Query Voltage Specification                       chr17040E_QueryVoltageSpecification

Query Current Specification                       chr17040E_QueryCurrentSpecification

Query Power Specification                         chr17040E_QueryPowerSpecification

Query Current Slew Rate Specification             chr17040E_QueryCurrentSlewRateSpecification

Battery Simulator

Configure Battery Total Capacity                     chr17040E_ConfigureBatteryTotalCapacity

Configure Battery Charge State Voltage               chr17040E_ConfigureBatteryChargeStateVoltage

Configure Battery Initial Capacity Mode              chr17040E_ConfigureBatteryInitialCapacityMode

Configure Battery Initial Voltage Mode               chr17040E_ConfigureBatteryInitialVoltageMode

Configure Battery Over Current Protection            chr17040E_ConfigureBatteryOverCurrentProtection

Configure Battery Internal Resistance                chr17040E_ConfigureBatteryInternalResistance

Configure Battery SOC Warning Limit                  chr17040E_ConfigureBatterySOCWarningLimit

Configure Battery Voltage Warning Limit              chr17040E_ConfigureBatteryVoltageWarningLimit

Configure Battery Efficiency                         chr17040E_ConfigureBatteryEfficiency

Configure Battery SOC Protection Limit               chr17040E_ConfigureBatterySOCProtectionLimit

Configure Battery Voltage Protection Limit           chr17040E_ConfigureBatteryVoltageProtectionLimit

Configure Battery Voltage Slew Rate                  chr17040E_ConfigureBatteryVoltageSlewRate

Upload Battery Voltage Curve                         chr17040E_UploadBatteryVoltageCurve

Upload Battery Capacity Curve                        chr17040E_UploadBatteryCapacityCurve

Upload Battery Discharge DCR Curve                   chr17040E_UploadBatteryDischargeDCRCurve

Upload Battery Charge DCR Curve                      chr17040E_UploadBatteryChargeDCRCurve

Specification Query

Query Battery Voltage Specification               chr17040E_QueryBatteryVoltageSpecification

Query Battery Current Specification               chr17040E_QueryBatteryCurrentSpecification

Query Battery Resistance Specification            chr17040E_QueryBatteryResistanceSpecification

Query Battery Voltage Protection Specification    chr17040E_QueryBatteryVoltageProtectionSpecification

Query Battery Voltage Slew Rate Specification     chr17040E_QueryBatteryVoltageSlewRateSpecification

Set/Get/Check Attribute

Set Attribute

Set Attribute ViInt32                             chr17040E_SetAttributeViInt32

Set Attribute ViReal64                            chr17040E_SetAttributeViReal64

Set Attribute ViString                            chr17040E_SetAttributeViString

Set Attribute ViBoolean                           chr17040E_SetAttributeViBoolean

Set Attribute ViSession                           chr17040E_SetAttributeViSession

Get Attribute

Get Attribute ViInt32                             chr17040E_GetAttributeViInt32

Get Attribute ViReal64                            chr17040E_GetAttributeViReal64

Get Attribute ViString                            chr17040E_GetAttributeViString

Get Attribute ViBoolean                           chr17040E_GetAttributeViBoolean

Get Attribute ViSession                           chr17040E_GetAttributeViSession

Check Attribute

Check Attribute ViInt32                           chr17040E_CheckAttributeViInt32

Check Attribute ViReal64                          chr17040E_CheckAttributeViReal64

Check Attribute ViString                          chr17040E_CheckAttributeViString

Check Attribute ViBoolean                         chr17040E_CheckAttributeViBoolean

Check Attribute ViSession                         chr17040E_CheckAttributeViSession

Action/Status Functions

Set Output State                                        chr17040E_SetOutputState

Set Battery Output State                                chr17040E_SetBatteryOutputState

Set Timer State                                         chr17040E_SetTimerState

Reset Timer                                             chr17040E_ResetTimer

Clear Protection                                        chr17040E_ClearProtection

Data Functions

Initiate Measurement                                    chr17040E_InitiateMeasurement

Fetch Operation Status                                  chr17040E_FetchOperationStatus

Fetch Operation Mode                                    chr17040E_FetchOperationMode

Fetch Time                                              chr17040E_FetchTime

Fetch Temperature Array                                 chr17040E_FetchTemeratureArray

Fetch Measurement                                       chr17040E_FetchMeasurement

Fetch Alarm Status                                      chr17040E_FetchAlarmStatus

Fetch Error Array                                       chr17040E_FetchErrorArray

System

Write Bytes                                             chr17040E_viWrite

Read Bytes                                              chr17040E_viRead

Utility Functions

Reset                                                   chr17040E_reset

Reset With Defaults                                     chr17040E_ResetWithDefaults

Disable                                                 chr17040E_Disable

Self-Test                                               chr17040E_self_test

Revision Query                                          chr17040E_revision_query

Error-Query                                             chr17040E_error_query

Error Message                                           chr17040E_error_message

Invalidate All Attributes                               chr17040E_InvalidateAllAttributes

Helper

Get Channel Name                                     chr17040E_GetChannelName

Is Not A Number                                      chr17040E_IsNotANumber

Error Info

Get Error                                            chr17040E_GetError

Clear Error                                          chr17040E_ClearError

Coercion Info

Get Next Coercion Record                             chr17040E_GetNextCoercionRecord

Interchangeability Info

Get Next Interchange Warning                         chr17040E_GetNextInterchangeWarning

Clear Interchange Warnings                           chr17040E_ClearInterchangeWarnings

Reset Interchange Check                              chr17040E_ResetInterchangeCheck

Locking

Lock Session                                         chr17040E_LockSession

Unlock Session                                       chr17040E_UnlockSession

Close                                                      chr17040E_close

Obsolete

Instrument I/O

Write Instrument Data                                chr17040E_WriteInstrData

Read Instrument Data                                 chr17040E_ReadInstrData

________________________________________________________________________________

Chroma 17040E

This instrument driver contains programming support for the Chroma

19301A.  This driver has all the functions that IVI and VXIplug&play

require.

Note:  This driver requires the VISA and IVI libraries.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

chr17040E_CheckAttributeViBoolean

ViStatus chr17040E_CheckAttributeViBoolean (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViBoolean attributeValue);

Purpose

This function checks the validity of a value you specify for a ViBoolean

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViBoolean

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViBoolean are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViBoolean

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_CheckAttributeViInt32

ViStatus chr17040E_CheckAttributeViInt32 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViInt32 attributeValue);

Purpose

This function checks the validity of a value you specify for a ViInt32

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViInt32 type.

If you choose to see all IVI attributes, the data types appear

to the right of the attribute names in the list box.

Attributes with data types other than ViInt32 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViInt32

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_CheckAttributeViReal64

ViStatus chr17040E_CheckAttributeViReal64 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViReal64 attributeValue);

Purpose

This function checks the validity of a value you specify for a ViReal64

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViReal64

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViReal64 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViReal64

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_CheckAttributeViSession

ViStatus chr17040E_CheckAttributeViSession (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViSession attributeValue);

Purpose

This function checks the validity of a value you specify for a ViSession

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViSession

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViSession are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViSession

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_CheckAttributeViString

ViStatus chr17040E_CheckAttributeViString (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViChar _VI_FAR attributeValue[]);

Purpose

This function checks the validity of a value you specify for a ViString

attribute.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to check the attribute value. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViString

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViString are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViChar[]

Pass the value which you want to verify as a valid value for the

attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ClearError

ViStatus chr17040E_ClearError (ViSession instrumentHandle);

Purpose

This function clears the error code and error description for the IVI

session. If the user specifies a valid IVI session for the

instrument_handle parameter, this function clears the error information

for the session. If the user passes VI_NULL for the Vi parameter, this

function clears the error information for the current execution thread.

If the Vi parameter is an invalid session, the function does nothing and

returns an error.

The function clears the error code by setting it to VI_SUCCESS.  If the

error description string is non-NULL, the function de-allocates the error

description string and sets the address to VI_NULL.

Maintaining the error information separately for each thread is useful if

the user does not have a session handle to pass to the chr17040E_GetError

function, which occurs when a call to chr17040E_init or

chr17040E_InitWithOptions fails.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ClearInterchangeWarnings

ViStatus chr17040E_ClearInterchangeWarnings

(ViSession instrumentHandle);

Purpose

This function clears the list of current interchange warnings.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ClearProtection

ViStatus chr17040E_ClearProtection (ViSession instrumentHandle,

ViConstString channelName);

Purpose

This function clears the occurred protection.

Remote-control command:

PROTection:CLEar

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_close

ViStatus chr17040E_close (ViSession instrumentHandle);

Purpose

This function performs the following operations:

- Closes the instrument I/O session.

- Destroys the instrument driver session and all of its attributes.

- Deallocates any memory resources the driver uses.

Notes:

(1) You must unlock the session before calling chr17040E_close.

(2) After calling chr17040E_close, you cannot use the instrument driver

again until you call chr17040E_init or chr17040E_InitWithOptions.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryChargeStateVoltage

ViStatus chr17040E_ConfigureBatteryChargeStateVoltage

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 fullStateVoltage, ViReal64 emptyStateVoltage);

Purpose

This function configures the voltage at 0% SOC and the voltage at 100%

SOC.

Equivalent parameter:

VH - Full State Voltage (at 100% SOC)

VL - Empty State Voltage (at 0% SOC)

Remote-control command:

BATTery:VH - Full State Voltage

BATTery:VL - Empty State Voltage

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

fullStateVoltage

Variable Type       ViReal64

Sets the voltage at 100% SOC.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 500.00 V

emptyStateVoltage

Variable Type       ViReal64

Sets the voltage at 0% SOC.

Valid Values: <Spec. Min> ~ <Full State Voltage> V

Default Value: 10.00 V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryEfficiency

ViStatus chr17040E_ConfigureBatteryEfficiency

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 chargeEfficiency, ViReal64 dischargeEfficiency);

Purpose

This function configures the charge/discharge efficiency.

Equivalent parameter:

EFF_CHG - Charge Efficiency

EFF_DSG - Discharge Efficiency

Remote-control command:

BATTery:EFFCHG - Charge Efficiency

BATTery:EFFDSG - Discharge Efficiency

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

chargeEfficiency

Variable Type       ViReal64

Sets the charge efficiency.

Valid Values: 0 ~ 100.0 %

Default Value: 100.0 %

dischargeEfficiency

Variable Type       ViReal64

Sets the discharge efficiency.

Valid Values: 0 ~ 100.0 %

Default Value: 100.0 %

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryInitialCapacityMode

ViStatus chr17040E_ConfigureBatteryInitialCapacityMode

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 initialCapacity);

Purpose

This function sets the battery simulator to initial capacity mode and

configures the initial value of capacity.

Remote-control command:

BATTery:INITial - Initial Mode

BATTery:INITial:CAP - Initial Capacity Value

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

initialCapacity

Variable Type       ViReal64

Sets the initial value of capacity.

Valid Values: 0 ~ 100.0 %

Default Value: 0 %

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryInitialVoltageMode

ViStatus chr17040E_ConfigureBatteryInitialVoltageMode

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 initialVoltage);

Purpose

This function sets the battery simulator to initial voltage mode and

configures the initial voltage level.

Remote-control command:

BATTery:INITial - Initial Mode

BATTery:INITial:VOL - Initial Voltage Level

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

initialVoltage

Variable Type       ViReal64

Sets the initial voltage level.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryBatteryVoltageSpecification function to obtain

the maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryInternalResistance

ViStatus chr17040E_ConfigureBatteryInternalResistance

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 internalResistance);

Purpose

This function configures the internal resistance of the battery

simulator.

Equivalent parameter:

ESR

Remote-control command:

BATTery:ESR

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

internalResistance

Variable Type       ViReal64

Sets the internal resistance of battery.

Valid Values: <Spec. Min> ~ <Spec. Max> ohm

Default Value: 1.000 ohm

Note:

Call chr17040E_QueryBatteryResistanceSpecification function to obtain

the maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryOverCurrentProtection

ViStatus chr17040E_ConfigureBatteryOverCurrentProtection

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 OCPLimit);

Purpose

This function configures the over current protection limit.

Equivalent parameter:

OCP

Remote-control command:

BATTery:OCP

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

OCPLimit

Variable Type       ViReal64

Sets the OCP limit.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryBatteryCurrentSpecification function to obtain

the maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatterySOCProtectionLimit

ViStatus chr17040E_ConfigureBatterySOCProtectionLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 overSOCLimit, ViReal64 underSOCLimit);

Purpose

This function configures the over/under SOC (State of Charge) protection

limit. The battery simulator disables the output when the SOC reaches the

given limits.

Remote-control command:

BATTery:BOH - Over SOC Limit

BATTery:BOL - Under SOC Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

overSOCLimit

Variable Type       ViReal64

Sets the over SOC protection limit. That is the BOH parameter of the

battery simulator. The battery simulator disables the output when the

SOC is greater than or equal to the value of this input.

Valid Values: 0 ~ 100.0 %

Default Value: 100.0 %

underSOCLimit

Variable Type       ViReal64

Sets the under SOC protection limit. That is the BOL parameter of the

battery simulator. The battery simulator disables the output when the

SOC is less than or equal to the value of this input.

Valid Values: 0 ~ <Over SOC Limit> %

Default Value: 0 %

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatterySOCWarningLimit

ViStatus chr17040E_ConfigureBatterySOCWarningLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 overSOCLimit, ViReal64 underSOCLimit);

Purpose

This function configures the over/under SOC (State Of Charge) warning

limit. The battery simulator shows warning but keeps output when the SOC

reaches the given limits.

Equivalent parameter:

BCH - Over SOC Limit

BCL - Under SOC Limit

Remote-control command:

BATTery:BCH - Over SOC Limit

BATTery:BCL - Under SOC Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

overSOCLimit

Variable Type       ViReal64

Sets the over SOC warning limit. The battery simulator shows warning

but keeps output when the SOC is greater than or equal to the value

of this input.

Valid Values: 0 ~ 100.0 %

Default Value: 100.0 %

underSOCLimit

Variable Type       ViReal64

Sets the under SOC warning limit. The battery simulator shows warning

but keeps output when the SOC is less than or equal to the value of

this input.

Valid Values: 0 ~ <Over SOC Limit> %

Default Value: 0 %

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryTotalCapacity

ViStatus chr17040E_ConfigureBatteryTotalCapacity

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 totalCapacity);

Purpose

This function configures the total capacity of battery.

Remote-control command:

BATTery:CAP

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

totalCapacity

Variable Type       ViReal64

Sets the total capacity

Valid Values: 0 ~ 9999.99 Ah

Default Value: 9999.99 Ah

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryVoltageProtectionLimit

ViStatus chr17040E_ConfigureBatteryVoltageProtectionLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 overVoltageLimit, ViReal64 underVoltageLimit);

Purpose

This function configures the over/under voltage protection limit. The

battery simulator disables the output when the voltage reaches the given

limits.

Remote-control command:

BATTery:VOH - Over Voltage Limit

BATTery:VOL - Under Voltage Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

overVoltageLimit

Variable Type       ViReal64

Sets the over voltage protection limit. That is the BVH parameter of

battery simulator. The battery simulator disables the output when the

voltage is greater than or equal to the value of this input.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 500.00 V

Note:

Call chr17040E_QueryBatteryVoltageSpecification function to obtain

the maximum and minimum.

underVoltageLimit

Variable Type       ViReal64

Sets the under voltage protection limit. That is the BVL parameter of

the battery simulator. The battery simulator disables the output when

the voltage is less than or equal to the value of this input.

Valid Values: <Spec. Min> ~ <Over Voltage Limit> V

Default Value: 100.00 V

Note:

Call chr17040E_QueryBatteryVoltageSpecification function to obtain

the maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryVoltageSlewRate

ViStatus chr17040E_ConfigureBatteryVoltageSlewRate

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageSlewRate);

Purpose

This function configures the voltage slew rate of battery simulator.

Remote-control command:

BATTery:VOLT:SLEW

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageSlewRate

Variable Type       ViReal64

Sets the voltage slew rate.

Valid Values: <Spec. Min> ~ <Spec. Max> V/ms

Default Value: 1.000 V/ms

Note:

Call chr17040E_QueryBatteryVoltageSlewRateSpecification function to

obtain the maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureBatteryVoltageWarningLimit

ViStatus chr17040E_ConfigureBatteryVoltageWarningLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 overVoltageLimit, ViReal64 underVoltageLimit);

Purpose

This function configures the over/under voltage warning limit. The

battery simulator shows warning but keeps output when the voltage reaches

the given limits.

Equivalent parameter:

BVH - Over Voltage Limit

BVL - Under Voltage Limit

Remote-control command:

BATTery:BVH - Over Voltage Limit

BATTery:BVL - Under Voltage Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

overVoltageLimit

Variable Type       ViReal64

Sets the over voltage warning limit. That is the BVH parameter of

battery simulator. The batter simulator shows warning but keeps

output when the voltage is greater than or equal to the value of this

input.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 500.00 V

Note:

Call chr17040E_QueryBatteryVoltageSpecification function to obtain

the maximum and minimum.

underVoltageLimit

Variable Type       ViReal64

Sets the under voltage warning limit. That is the BVL parameter of

battery simulator. The batter simulator shows warning but keeps

output when the voltage is less than or equal to the value of this

input.

Valid Values: <Spec. Min> ~ <Over Voltage Limit> V

Default Value: 100.00 V

Note:

Call chr17040E_QueryBatteryVoltageSpecification function to obtain

the maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCCChargeCurrentLevel

ViStatus chr17040E_ConfigureCCChargeCurrentLevel

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 currentLevel);

Purpose

This function configures the current level in constant current charge

mode. This function sets the cutoff current to 0.

Remote-control command:

SOURce:CURR - Current Level

SOURce:CURRent:CUTOFF - Cutoff Current

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentLevel

Variable Type       ViReal64

Sets the current level.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCCChargePowerLimit

ViStatus chr17040E_ConfigureCCChargePowerLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 powerLimit);

Purpose

This function configures the power limit in constant current charge mode.

Remote-control command:

SOURce:POW - Power Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

powerLimit

Variable Type       ViReal64

Sets the power limit.

Valid Values: <Spec. Min> ~ <Spec. Max> W

Default Value: 0 W

Note:

Call chr17040E_QueryPowerSpecification function to obtain the maximum

and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCCChargeVoltageCondition

ViStatus chr17040E_ConfigureCCChargeVoltageCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageLevel, ViReal64 cutoffVoltage);

Purpose

This function configures the voltage level and cutoff voltage in constant

current charge mode.

Remote-control command:

SOURce:VOLT - Voltage Level

SOURce:VOLTage:CUTOFF - Cutoff Voltage

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageLevel

Variable Type       ViReal64

Sets the voltage level. This value must be greater than or equal to

the cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

cutoffVoltage

Variable Type       ViReal64

Sets the cutoff voltage for charging.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCCDischargeCurrentLevel

ViStatus chr17040E_ConfigureCCDischargeCurrentLevel

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 currentLevel);

Purpose

This function configures the current level in constant current discharge

mode. This function sets the cutoff current to 0.

Remote-control command:

SOURce:CURR - Current Level

SOURce:CURRent:CUTOFF - Cutoff Current

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentLevel

Variable Type       ViReal64

Sets the current level.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCCDischargePowerLimit

ViStatus chr17040E_ConfigureCCDischargePowerLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 powerLimit);

Purpose

This function configures the power limit in constant current discharge

mode.

Remote-control command:

SOURce:POW - Power Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

powerLimit

Variable Type       ViReal64

Sets the power limit.

Valid Values: <Spec. Min> ~ <Spec. Max> W

Default Value: 0 W

Note:

Call chr17040E_QueryPowerSpecification function to obtain the maximum

and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCCDischargeVoltageCondition

ViStatus chr17040E_ConfigureCCDischargeVoltageCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageLevel, ViReal64 cutoffVoltage);

Purpose

This function configures the voltage level and the cutoff voltage in

constant current discharge mode.

Remote-control command:

SOURce:VOLT - Voltage Level

SOURce:VOLTage:CUTOFF - Cutoff Voltage

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageLevel

Variable Type       ViReal64

Sets the voltage level. This value must be less than or equal to the

cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

cutoffVoltage

Variable Type       ViReal64

Sets the cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCPChargeCurrentCondition

ViStatus chr17040E_ConfigureCPChargeCurrentCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 currentLimit);

Purpose

This function configures the current limit in constant power charge mode.

This function sets the cutoff current to 0.

Remote-control command:

SOURce:CURR - Current Limit

SOURce:CURRent:CUTOFF - Cutoff Current

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentLimit

Variable Type       ViReal64

Sets the current limit.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCPChargePowerLevel

ViStatus chr17040E_ConfigureCPChargePowerLevel

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 powerLevel);

Purpose

This function configures the power level in constant power charge mode.

Remote-control command:

SOURce:POW - Power Level

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

powerLevel

Variable Type       ViReal64

Sets the power level.

Valid Values: <Spec. Min> ~ <Spec. Max> W

Default Value: 0 W

Note:

Call chr17040E_QueryPowerSpecification function to obtain the maximum

and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCPChargeVoltageCondition

ViStatus chr17040E_ConfigureCPChargeVoltageCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageLevel, ViReal64 cutoffVoltage);

Purpose

This function configures the voltage level and cutoff voltage in constant

power charge mode.

Remote-control command:

SOURce:VOLT - Voltage Level

SOURce:VOLTage:SLEW - Cutoff Voltage

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageLevel

Variable Type       ViReal64

Sets the voltage level. This value must be greater than or equal to

the cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

cutoffVoltage

Variable Type       ViReal64

Sets the cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCPDischargeCurrentCondition

ViStatus chr17040E_ConfigureCPDischargeCurrentCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 currentLimit);

Purpose

This function configures the current limit in constant power discharge

mode. This function sets the cutoff current to 0.

Remote-control command:

SOURce:CURR - Current Limit

SOURce:CURRent:CUTOFF - Cutoff Current

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentLimit

Variable Type       ViReal64

Sets the current limit.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCPDischargePowerLevel

ViStatus chr17040E_ConfigureCPDischargePowerLevel

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 powerLevel);

Purpose

This function configures the power level in constant power discharge

mode.

Remote-control command:

SOURce:POW - Power Level

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

powerLevel

Variable Type       ViReal64

Sets the power level.

Valid Values: <Spec. Min> ~ <Spec. Max> W

Default Value: 0 W

Note:

Call chr17040E_QueryPowerSpecification function to obtain the maximum

and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCPDischargeVoltageCondition

ViStatus chr17040E_ConfigureCPDischargeVoltageCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageLevel, ViReal64 cutoffVoltage);

Purpose

This function configures the voltage level and cutoff voltage in constant

power discharge mode.

Remote-control command:

SOURce:VOLT - Voltage Level

SOURce:VOLTage:SLEW - Cutoff Voltage

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageLevel

Variable Type       ViReal64

Sets the voltage level. This value must be less than or equal to the

cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

cutoffVoltage

Variable Type       ViReal64

Sets the cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCurrentSlewRate

ViStatus chr17040E_ConfigureCurrentSlewRate (ViSession instrumentHandle,

ViConstString channelName,

ViReal64 currentSlewRate);

Purpose

This function configures the current slew rate.

Remote-control command:

SOURce:CURRent:SLEW

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentSlewRate

Variable Type       ViReal64

Sets the current slew rate.

Valid Values: <Spec. Min> ~ <Spec. Max> A/ms

Default Value: 0 A/ms

Note:

Call chr17040E_QueryCurrentSlewRateSpecification function to obtain

the maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCutoffTime

ViStatus chr17040E_ConfigureCutoffTime (ViSession instrumentHandle,

ViConstString channelName,

ViReal64 cutoffTime);

Purpose

This function configures the cutoff time. The controller will disable

output when the output time is greater than the cutoff time. You can

input CHR17040E_VAL_INFINITE_CUTOFF_TIME to set a infinite cutoff time.

Remote-control command:

SOURce:TIME:CUTOFF

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

cutoffTime

Variable Type       ViReal64

Sets the cutoff time.

Valid Values:

CHR17040E_VAL_INFINITE_CUTOFF_TIME - infinite time

1 ~ 65535 sec - finite time

Default Value: 1

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVChargeCurrentCondition

ViStatus chr17040E_ConfigureCVChargeCurrentCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 currentLimit, ViReal64 cutoffCurrent);

Purpose

This function configures the current limit and the cutoff current in

constant voltage charge mode.

Remote-control command:

SOURce:CURR - Current Limit

SOURce:CURRent:CUTOFF - Cutoff Current

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentLimit

Variable Type       ViReal64

Sets the current limit. This value must be greater than or equal to

the cutoff current.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

cutoffCurrent

Variable Type       ViReal64

Sets the cutoff current.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVChargePowerLimit

ViStatus chr17040E_ConfigureCVChargePowerLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 powerLimit);

Purpose

This function configures the power limit in constant voltage charge mode.

Remote-control command:

SOURce:POW - Power Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

powerLimit

Variable Type       ViReal64

Sets the power limit.

Valid Values: <Spec. Min> ~ <Spec. Max> W

Default Value: 0 W

Note:

Call chr17040E_QueryPowerSpecification function to obtain the maximum

and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVChargeVoltageLevel

ViStatus chr17040E_ConfigureCVChargeVoltageLevel

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageLevel);

Purpose

This function configures the voltage level in constant voltage charge

mode. This function sets the cutoff voltage to 0.

Remote-control command:

SOURce:VOLT - Voltage Level

SOURce:VOLTage:CUTOFF - Cutoff Voltage

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageLevel

Variable Type       ViReal64

Sets the voltage level.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVDischargeCurrentCondition

ViStatus chr17040E_ConfigureCVDischargeCurrentCondition

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 currentLimit, ViReal64 cutoffCurrent);

Purpose

This function configures the current limit and the cutoff current in

constant voltage discharge mode.

Remote-control command:

SOURce:CURR - Current Limit

SOURce:CURRent:CUTOFF - Cutoff Current

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentLimit

Variable Type       ViReal64

Sets the current limit. This value must be greater than or equal to

the cutoff current.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

cutoffCurrent

Variable Type       ViReal64

Sets the cutoff current.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVDischargePowerLimit

ViStatus chr17040E_ConfigureCVDischargePowerLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 powerLimit);

Purpose

This function configures the power limit in constant voltage discharge

mode.

Remote-control command:

SOURce:POW - Power Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

powerLimit

Variable Type       ViReal64

Sets the power limit.

Valid Values: <Spec. Min> ~ <Spec. Max> W

Default Value: 0 W

Note:

Call chr17040E_QueryPowerSpecification function to obtain the maximum

and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVDischargeVoltageLevel

ViStatus chr17040E_ConfigureCVDischargeVoltageLevel

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageLevel);

Purpose

This function configures the voltage level in constant current discharge

mode. This function sets the cutoff voltage to 0.

Remote-control command:

SOURce:VOLT - Voltage Level

SOURce:VOLTage:CUTOFF - Cutoff Voltage

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageLevel

Variable Type       ViReal64

Sets the voltage level.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVSourceCurrentLimit

ViStatus chr17040E_ConfigureCVSourceCurrentLimit

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 currentLimit);

Purpose

This function configures the current limit.

Remote-control command:

SOURce:CURR - Current Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

currentLimit

Variable Type       ViReal64

Sets the current limit.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Note:

Call chr17040E_QueryCurrentSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureCVSourceVoltageLevel

ViStatus chr17040E_ConfigureCVSourceVoltageLevel

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 voltageLevel);

Purpose

This function configures the voltage level.

Remote-control command:

SOURce:VOLT - Voltage Level

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

voltageLevel

Variable Type       ViReal64

Sets the voltage level.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Note:

Call chr17040E_QueryVoltageSpecification function to obtain the

maximum and minimum.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureNegativeDeltaCurrentProtection

ViStatus chr17040E_ConfigureNegativeDeltaCurrentProtection

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 UDCPLimit);

Purpose

This function configures the negative delta current protection. The

controller disables the output when the negative current difference is

greater than or equal to the value of UDCP Limit.

Equivalent parameter:

UDCP

Remote-control command:

[:SOURce]:CURRent:PROTection:DELTa:UNDer[:LEVel]

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

UDCPLimit

Variable Type       ViReal64

Sets the negative delta current protection limit. The controller

disables the output when the negative current difference is greater

than or equal to the value of this input.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureNegativeDeltaVoltageProtection

ViStatus chr17040E_ConfigureNegativeDeltaVoltageProtection

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 UDVPLimit);

Purpose

This function configures the negative delta voltage protection. The

controller disables the output when the negative voltage difference is

greater than or equal to the value of UDVP Limit.

Equivalent parameter:

UDVP

Remote-control command:

[:SOURce]:VOLTage:PROTection:DELTa:UNDer[:LEVel]

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

UDVPLimit

Variable Type       ViReal64

Sets the negative delta voltage protection limit. The controller

disables the output when the negative voltage difference is greater

than or equal to the value of this input.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureOverCapacityProtection

ViStatus chr17040E_ConfigureOverCapacityProtection

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 OQPLimit);

Purpose

This function configures the over capacity protection. The controller

disables the output when the capacity is greater than or equal to the

value of OQP Limit.

Equivalent parameter:

OQP

Remote-control command:

AH:PROTection

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

OQPLimit

Variable Type       ViReal64

Sets the over capacity protection limit. The controller disables the

output when the capacity is greater than or equal to the value of

this input.

Valid Values: 0 ~ <+Inf> Ah

Default Value: 0 Ah

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureOverCurrentProtection

ViStatus chr17040E_ConfigureOverCurrentProtection

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 OCPLimit);

Purpose

This function configures the over current protection (OCP). The

controller disables the output when the current is greater than or equal

to the value of OCP Limit.

Remote-control command:

[:SOURce]:CURRent:PROTection[:LEVel]

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

OCPLimit

Variable Type       ViReal64

Sets the over current protection limit. The controller disables the

output when the current is greater than or equal to the value of this

input.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureOverVoltageProtection

ViStatus chr17040E_ConfigureOverVoltageProtection

(ViSession instrumentHandle, ViConstString channelName,

ViBoolean OVPEnable, ViReal64 OVPLimit);

Purpose

This function configures the over voltage protection (OVP). The

controller disables the output when the voltage is greater than or equal

to the value of OVP Limit.

Remote-control command:

[:SOURce]:VOLTage:PROTection:STATe - OVP Enable

[:SOURce]:VOLTage:PROTection[:LEVel] - OVP Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

OVPEnable

Variable Type       ViBoolean

Enables the over voltage protection.

Valid Values:

VI_TRUE - Enable Protection

VI_FALSE - Disable Protection

Default Value: VI_FALSE

OVPLimit

Variable Type       ViReal64

Sets the over voltage protection limit. The controller disables the

output when the voltage is greater than or equal to the value of this

input.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigurePositiveDeltaCurrentProtection

ViStatus chr17040E_ConfigurePositiveDeltaCurrentProtection

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 ODCPLimit);

Purpose

This function configures the positive delta current protection. The

controller disables the output when the positive current difference is

greater than or equal to the value of ODCP Limit.

Equivalent parameter:

ODCP

Remote-control command:

[:SOURce]:CURRent:PROTection:DELTa[:LEVel]

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

ODCPLimit

Variable Type       ViReal64

Sets the positive delta current protection limit. The controller

disables the output when the positive current difference is greater

than or equal to the value of this input.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigurePositiveDeltaVoltageProtection

ViStatus chr17040E_ConfigurePositiveDeltaVoltageProtection

(ViSession instrumentHandle, ViConstString channelName,

ViReal64 ODVPLimit);

Purpose

This function configures the positive delta voltage protection. The

controller disables the output when the positive voltage difference is

greater than or equal to the value of ODVP Limit.

Equivalent parameter:

ODVP

Remote-control command:

[:SOURce]:VOLTage:PROTection:DELTa[:LEVel]

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

ODVPLimit

Variable Type       ViReal64

Sets the positive delta voltage protection limit. The controller

disables the output when the positive voltage difference is greater

than or equal to the value of this input.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureSourceType

ViStatus chr17040E_ConfigureSourceType (ViSession instrumentHandle,

ViConstString channelName,

ViInt32 sourceType);

Purpose

This function configures the output source type.

Remote-control command:

SOURce:MODe

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

sourceType

Variable Type       ViInt32

Sets the output source type.

Valid Values:

CHR17040E_VAL_CC_CHARGE - CC Charge

CHR17040E_VAL_CV_CHARGE - CV Charge

CHR17040E_VAL_CP_CHARGE - CP Charge

CHR17040E_VAL_CC_DISCHARGE - CC Discharge

CHR17040E_VAL_CV_DISCHARGE - CV Discharge

CHR17040E_VAL_CP_DISCHARGE - CP Discharge

CHR17040E_VAL_CV_SOURCE - CV Source

CHR17040E_VAL_REST - Reset

Default Value: CHR17040E_VAL_CC_CHARGE

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ConfigureUnderVoltageProtection

ViStatus chr17040E_ConfigureUnderVoltageProtection

(ViSession instrumentHandle, ViConstString channelName,

ViBoolean UVPEnable, ViReal64 UVPLimit);

Purpose

This function configures the under voltage protection (UVP). The

controller disables the output when the voltage is less than or equal to

the value of UVP Limit.

Remote-control command:

[:SOURce]:VOLTage:PROTection:UNDer:STATe - UVP Enable

[:SOURce]:VOLTage:PROTection:UNDer[:LEVel] - UVP Limit

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

UVPEnable

Variable Type       ViBoolean

Enables the under voltage protection.

Valid Values:

VI_TRUE - Enable Protection

VI_FALSE - Disable Protection

Default Value: VI_FALSE

UVPLimit

Variable Type       ViReal64

Sets the under voltage protection limit. The controller disables the

output when the voltage is less than or equal to the value of this

input.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_Disable

ViStatus chr17040E_Disable (ViSession instrumentHandle);

Purpose

This function places the instrument in a quiescent state where it has

minimal or no impact on the system to which it is connected.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_error_message

ViStatus chr17040E_error_message (ViSession instrumentHandle,

ViStatus errorCode,

ViChar _VI_FAR errorMessage[]);

Purpose

This function converts a status code returned by an instrument driver

function into a user-readable string.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

You can pass VI_NULL for this parameter.  This is useful when one of

the initialize functions fail.

Default Value:  VI_NULL

errorCode

Variable Type       ViStatus

Pass the Status parameter that is returned from any of the instrument

driver functions.

Default Value:  0  (VI_SUCCESS)

errorMessage

Variable Type       ViChar[]

Returns the user-readable message string that corresponds to the

status code you specify.

You must pass a ViChar array with at least 256 bytes.

Return Value

Reports the status of this operation.

This function can return only three possible status codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0085  Unknown status code (warning).

BFFF000A  Invalid parameter (Error Message buffer is VI_NULL).

________________________________________________________________________________

chr17040E_error_query

ViStatus chr17040E_error_query (ViSession instrumentHandle,

ViPInt32 errorCode,

ViChar _VI_FAR errorMessage[]);

Purpose

This function reads an error code and a message from the instrument's

error queue.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

errorCode

Variable Type       ViInt32 (passed by reference)

Returns the error code read from the instrument's error queue.

errorMessage

Variable Type       ViChar[]

Returns the error message string read from the instrument's error

message queue.

You must pass a ViChar array with at least 256 bytes.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_FetchAlarmStatus

ViStatus chr17040E_FetchAlarmStatus (ViSession instrumentHandle,

ViConstString channelName,

ViPInt32 statusCode);

Purpose

This function fetches the alarm status code from the driver cache.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

statusCode

Variable Type       ViInt32 (passed by reference)

Returns the alarm status code.

Meanings of status code:

Bit0: Over SOC warning, BCH

Bit1: Under SOC warning, BCL

Bit2: Over voltage warning, BVH

Bit3: Under voltage warning, BVL

Bit4: Test Debug Mode

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_FetchErrorArray

ViStatus chr17040E_FetchErrorArray (ViSession instrumentHandle,

ViConstString channelName,

ViInt32 _VI_FAR errorStatusArray[]);

Purpose

This function fetches the error statuses from the driver cache.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

errorStatusArray

Variable Type       ViInt32[]

Returns the error statuses.

Array Size: 3

Meanings of Error 1 status:

Bit0: BUS_OVP

Bit1: BUS_UVP

Bit2: OUT_OVP

Bit3: OUT_UVP

Bit4: BUS_OCP

Bit5: IL_OCP

Bit6: IL_UCP

Bit7: reserved

Bit8: reserved

Bit9: reserved

Bit10: PWM_FAIL

Bit11: FAN_FAIL

Bit12: OTP

Bit13: VO_NOT_EQUAL

Bit14: CONTACT_FAIL

Bit15: CALIBRATION_ERROR

Bit16: EMERGENCY_STOP

Bit17: AD_ERROR

Bit18: AD_OFFLINE

Bit19: BAT_BOH

Bit20: BAT_BOL

Bit21: BAT_VOH

Bit22: BAT_VOL

Bit23: temperature sensor error

Meanings of Error 2 status:

Reserved

Meanings of Error 3 status:

Bit0 ~ Bit15: not used

Bit16: CSU_DD_SYNC_ERR

Bit17: CSU_AD_SYNC_ERR

Bit18: CSU_IPC_TIMEOUT

Bit19: reserved

Bit20: CSU_DD_GROUP_ERR

Bit21: CSU_DD_TIMEOUT

Bit22: CSU_DD_COMM_TIMEOUT

Bit23: CSU_DD_SLAVE_ERR

Bit24: CSU_AD_COMM_TIMEOUT

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_FetchMeasurement

ViStatus chr17040E_FetchMeasurement (ViSession instrumentHandle,

ViConstString channelName,

ViInt32 measurementType,

ViPReal64 measurement);

Purpose

This function fetches the measured value of the given type from the

driver cache.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

measurementType

Variable Type       ViInt32

Inputs the type of measurement to fetch.

Valid Values:

CHR17040E_VAL_MEASURE_VOLTAGE - Voltage

CHR17040E_VAL_MEASURE_CURRENT - Current

CHR17040E_VAL_MEASURE_POWER - Power

CHR17040E_VAL_MEASURE_CAPACITY - Capacity

CHR17040E_VAL_MEASURE_WORK - Work

CHR17040E_VAL_MEASURE_DCIR - DCIR

Default Value: CHR17040E_VAL_MEASURE_VOLTAGE

measurement

Variable Type       ViReal64 (passed by reference)

Returns the measured value of given type.

Units of measurement types:

Voltage: V

Current: A

Power: W

Capacity: Ah

Work: kW-h

DCIR: ohm

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_FetchOperationMode

ViStatus chr17040E_FetchOperationMode (ViSession instrumentHandle,

ViConstString channelName,

ViPInt32 operationMode);

Purpose

This function fetches operation mode from the driver cache.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

operationMode

Variable Type       ViInt32 (passed by reference)

Returns the operation status.

Defined Values:

CHR17040E_VAL_OPERATION_MODE_STOP - Stop

CHR17040E_VAL_OPERATION_MODE_RUN - Run

CHR17040E_VAL_OPERATION_MODE_Pause - Pause

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_FetchOperationStatus

ViStatus chr17040E_FetchOperationStatus (ViSession instrumentHandle,

ViConstString channelName,

ViPInt32 operationStatus);

Purpose

This function fetches operation status from the driver cache.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

operationStatus

Variable Type       ViInt32 (passed by reference)

Returns the operation status.

Defined Values:

CHR17040E_VAL_STAUS_STOP_MODE - Stop Mode

CHR17040E_VAL_STAUS_SINGLE_MODE - Single Mode

CHR17040E_VAL_STAUS_PROGRAM_MODE - Program Mode

CHR17040E_VAL_STAUS_CALIBRATION_MODE - Calibration Mode

CHR17040E_VAL_STAUS_TEST_MODE - Test Mode

CHR17040E_VAL_STAUS_PAUSE_MODE - Pause Mode

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_FetchTemeratureArray

ViStatus chr17040E_FetchTemeratureArray (ViSession instrumentHandle,

ViConstString channelName,

ViReal64 _VI_FAR temeratureArray[]);

Purpose

This function fetches the temperatures from sensor 1 to sensor 8 from the

driver cache.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

temeratureArray

Variable Type       ViReal64[]

Returns the temperatures from sensor 1 to sensor 8.

Array Size: 8

Units: Degrees Celsius

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_FetchTime

ViStatus chr17040E_FetchTime (ViSession instrumentHandle,

ViConstString channelName,

ViPReal64 time);

Purpose

This function fetches the measurement time from the driver cache.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

time

Variable Type       ViReal64 (passed by reference)

Returns the measurement time.

Units: sec

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetAttributeViBoolean

ViStatus chr17040E_GetAttributeViBoolean (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPBoolean attributeValue);

Purpose

This function queries the value of a ViBoolean attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViBoolean

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViBoolean are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViBoolean (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViBoolean variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetAttributeViInt32

ViStatus chr17040E_GetAttributeViInt32 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPInt32 attributeValue);

Purpose

This function queries the value of a ViInt32 attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViInt32 type.

If you choose to see all IVI attributes, the data types appear

to the right of the attribute names in the list box.

Attributes with data types other than ViInt32 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViInt32 (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViInt32 variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetAttributeViReal64

ViStatus chr17040E_GetAttributeViReal64 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPReal64 attributeValue);

Purpose

This function queries the value of a ViReal64 attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViReal64

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViReal64 are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViReal64 (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViReal64 variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetAttributeViSession

ViStatus chr17040E_GetAttributeViSession (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViPSession attributeValue);

Purpose

This function queries the value of a ViSession attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViSession

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViSession are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViSession (passed by reference)

Returns the current value of the attribute.  Pass the address of a

ViSession variable.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetAttributeViString

ViStatus chr17040E_GetAttributeViString (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViInt32 bufferSize,

ViChar _VI_FAR attributeValue[]);

Purpose

This function queries the value of a ViString attribute.

You can use this function to get the values of instrument- specific

attributes and inherent IVI attributes.  If the attribute represents an

instrument state, this function performs instrument I/O in the following

cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid.

You must provide a ViChar array to serve as a buffer for the value.  You

pass the number of bytes in the buffer as the Buffer Size parameter.  If

the current value of the attribute, including the terminating NUL byte,

is larger than the size you indicate in the Buffer Size parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an ASCII

NUL byte at the end of the buffer, and returns the buffer size you must

pass to get the entire value.  For example, if the value is "123456" and

the Buffer Size is 4, the function places "123" into the buffer and

returns 7.

If you want to call this function just to get the required buffer size,

you can pass 0 for the Buffer Size and VI_NULL for the Attribute Value

buffer.

If you want the function to fill in the buffer regardless of the   number

of bytes in the value, pass a negative number for the Buffer Size

parameter.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to obtain the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Help text is

shown for each attribute.  Select an attribute by

double-clicking on it or by selecting it and then pressing

<ENTER>.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViString

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViString are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Attribute Value parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Attribute Value buffer

parameter.

Default Value: 512

attributeValue

Variable Type       ViChar[]

The buffer in which the function returns the current value of the

attribute.  The buffer must be of type ViChar and have at least as

many bytes as indicated in the Buffer Size parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you specify 0 for the Buffer Size parameter, you can pass VI_NULL

for this parameter.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has named constants as valid values, you can view a

list of the constants by pressing <ENTER> on this control.

Select a value by double-clicking on it or by selecting it and

then pressing <ENTER>.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

If the current value of the return buffer, including the terminating

NUL byte, is larger than the size you indicate in the Buffer Size

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value.  For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetChannelName

ViStatus chr17040E_GetChannelName (ViSession instrumentHandle,

ViInt32 index, ViInt32 bufferSize,

ViChar _VI_FAR channelName[]);

Purpose

The function returns the channel name string that represent the index you

specified. By passing 0 for the buffer size, the caller can ascertain the

buffer size required to get the entire channel name string and then call

the function again with a sufficiently large buffer.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

index

Variable Type       ViInt32

Inputs the 1-based index of channel.

Valid Range: 1 ~ Total Channel Number

Default Value: 1

bufferSize

Variable Type       ViInt32

Inputs the buffer size of result stage name string. By passing buffer

size 0 to obtain the required buffer size which including null char.

Default Value: 256

channelName

Variable Type       ViChar[]

Returns the channel name that represent the index you specified.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetError

ViStatus chr17040E_GetError (ViSession instrumentHandle, ViPStatus code,

ViInt32 bufferSize,

ViChar _VI_FAR description[]);

Purpose

This function retrieves and then clears the IVI error information for the

session or the current execution thread. One exception exists: If the

BufferSize parameter is 0, the function does not clear the error

information. By passing 0 for the buffer size, the caller can ascertain

the buffer size required to get the entire error description string and

then call the function again with a sufficiently large buffer.

If the user specifies a valid IVI session for the InstrumentHandle

parameter, Get Error retrieves and then clears the error information for

the session.  If the user passes VI_NULL for the InstrumentHandle

parameter, this function retrieves and then clears the error information

for the current execution thread.  If the InstrumentHandle parameter is

an invalid session, the function does nothing and returns an error.

Normally, the error information describes the first error that occurred

since the user last called chr17040E_GetError or chr17040E_ClearError.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

code

Variable Type       ViStatus (passed by reference)

Returns the error code for the session or execution thread.

If you pass 0 for the Buffer Size, you can pass VI_NULL for this

parameter.

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Description parameter.

If the error description, including the terminating NUL byte,

contains more bytes than you indicate in this parameter, the function

copies BufferSize - 1 bytes into the buffer, places an ASCII NUL byte

at the end of the buffer, and returns the buffer size you must pass

to get the entire value.  For example, if the value is "123456" and

the Buffer Size is 4, the function places "123" into the buffer and

returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Description buffer

parameter.

Default Value:  None

description

Variable Type       ViChar[]

Returns the error description for the IVI session or execution

thread.  If there is no description, the function returns an empty

string.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.  If the error description,

including the terminating NUL byte, contains more bytes than you

indicate with the Buffer Size parameter, the function copies Buffer

Size - 1 bytes into the buffer, places an ASCII NUL byte at the end

of the buffer, and returns the buffer size you must pass to get the

entire value.  For example, if the value is "123456" and the Buffer

Size is 4, the function places "123" into the buffer and returns 7.

If you pass 0 for the Buffer Size, you can pass VI_NULL for this

parameter.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

If the current value of the return buffer, including the terminating

NUL byte, is larger than the size you indicate in the Buffer Size

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value.  For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetNextCoercionRecord

ViStatus chr17040E_GetNextCoercionRecord (ViSession instrumentHandle,

ViInt32 bufferSize,

ViChar _VI_FAR coercionRecord[]);

Purpose

This function returns the coercion information associated with the IVI

session.  This function retrieves and clears the oldest instance in which

the instrument driver coerced a value you specified to another value.

If you set the CHR17040E_ATTR_RECORD_COERCIONS attribute to VI_TRUE, the

instrument driver keeps a list of all coercions it makes on ViInt32 or

ViReal64 values you pass to instrument driver functions.  You use this

function to retrieve information from that list.

If the next coercion record string, including the terminating NUL byte,

contains more bytes than you indicate in this parameter, the function

copies Buffer Size - 1 bytes into the buffer, places an ASCII NUL byte at

the end of the buffer, and returns the buffer size you must pass to get

the entire value.  For example, if the value is "123456" and the Buffer

Size is 4, the function places "123" into the buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Coercion Record buffer

parameter.

The function returns an empty string in the Coercion Record parameter if

no coercion records remain for the session.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init

function.  The handle identifies a particular instrument session.

Default Value:  None

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Coercion Record parameter.

If the next coercion record string, including the terminating NUL

byte, contains more bytes than you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Coercion Record buffer

parameter.

Default Value:  None

coercionRecord

Variable Type       ViChar[]

Returns the next coercion record for the IVI session.  If there are

no coercion records, the function returns an empty string.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.  If the next coercion record

string, including the terminating NUL byte, contains more bytes than

you indicate with the Buffer Size parameter, the function copies

Buffer Size - 1 bytes into the buffer, places an ASCII NUL byte at

the end of the buffer, and returns the buffer size you must pass to

get the entire value.  For example, if the value is "123456" and the

Buffer Size is 4, the function places "123" into the buffer and

returns 7.

This parameter returns an empty string if no coercion records remain

for the session.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

If the current value of the return buffer, including the terminating

NUL byte, is larger than the size you indicate in the Buffer Size

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value.  For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_GetNextInterchangeWarning

ViStatus chr17040E_GetNextInterchangeWarning (ViSession instrumentHandle,

ViInt32 bufferSize,

ViChar _VI_FAR interchangeWarning[]);

Purpose

This function returns the interchangeability warnings associated with the

IVI session. It retrieves and clears the oldest instance in which the

class driver recorded an interchangeability warning.  Interchangeability

warnings indicate that using your application with a different instrument

might cause different behavior. You use this function to retrieve

interchangeability warnings.

The driver performs interchangeability checking when the

CHR17040E_ATTR_INTERCHANGE_CHECK attribute is set to VI_TRUE.

The function returns an empty string in the Interchange Warning parameter

if no interchangeability warnings remain for the session.

In general, the instrument driver generates interchangeability warnings

when an attribute that affects the behavior of the instrument is in a

state that you did not specify.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Interchange Warning parameter.

If the next interchangeability warning string, including the

terminating NUL byte, contains more bytes than you indicate in this

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value. For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Interchange Warning

buffer parameter.

Default Value:  None

interchangeWarning

Variable Type       ViChar[]

Returns the next interchange warning for the IVI session. If there

are no interchange warnings, the function returns an empty  string.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the next

interchangeability warning string, including the terminating NUL

byte, contains more bytes than you indicate with the Buffer Size

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value.  For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

This parameter returns an empty string if no interchangeability

warnings remain for the session.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

If the current value of the return buffer, including the terminating

NUL byte, is larger than the size you indicate in the Buffer Size

parameter, the function copies Buffer Size - 1 bytes into the buffer,

places an ASCII NUL byte at the end of the buffer, and returns the

buffer size you must pass to get the entire value.  For example, if

the value is "123456" and the Buffer Size is 4, the function places

"123" into the buffer and returns 7.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_init

ViStatus chr17040E_init (ViRsrc resourceName, ViBoolean IDQuery,

ViBoolean resetDevice,

ViPSession instrumentHandle);

Purpose

This function performs the following initialization actions:

- Creates a new IVI instrument driver session.

- Opens a session to the specified device using the interface and address

you specify for the Resource Name parameter.

- If the ID Query parameter is set to VI_TRUE, this function queries the

instrument ID and checks that it is valid for this instrument driver.

- If the Reset parameter is set to VI_TRUE, this function resets the

instrument to a known state.

- Sends initialization commands to set the instrument to the state

necessary for the operation of the instrument driver.

- Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Note:  This function creates a new session each time you invoke it.

Although you can open more than one IVI session for the same resource, it

is best not to do so.  You can use the same session in multiple program

threads.  You can use the chr17040E_LockSession and

chr17040E_UnlockSession functions to protect sections of code that

require exclusive access to the resource.

Content from the IVI specifications reproduced with permission from the

IVI Foundation.

The IVI Foundation and its member companies make no warranty of any kind

with regard to this material, including, but not limited to, the implied

warranties of merchantability and fitness for a particular purpose. The

IVI Foundation and its member companies shall not be liable for errors

contained herein or for incidental or consequential damages in connection

with the furnishing, performance, or use of this material.

Parameter List

resourceName

Variable Type       ViRsrc

Pass the resource name of the device to initialize.

You can also pass the name of a driver session or logical name that

you configure with the IVI Configuration utility.  The driver session

identifies a specific device and specifies the initial settings for

the session.  A logical Name identifies a particular driver session.

Refer to the following table below for the exact grammar to use for

this parameter.  Optional fields are shown in square brackets ([]).

Syntax

------------------------------------------------------

GPIB[board]::<primary address>[::secondary address]::INSTR

VXI[board]::<logical address>::INSTR

GPIB-VXI[board]::<logical address>::INSTR

ASRL<port>::INSTR

<LogicalName>

<Driver Session>

If you do not specify a value for an optional field, the following

values are used:

Optional Field - Value

------------------------------------------------------

board - 0

secondary address - none (31)

The following table contains example valid values for this parameter.

"Valid Value" - Description

------------------------------------------------------

"GPIB::22::INSTR" - GPIB board 0, primary address 22 no

secondary address

"GPIB::22::5::INSTR" - GPIB board 0, primary address 22

secondary address 5

"GPIB1::22::5::INSTR" - GPIB board 1, primary address 22

secondary address 5

"VXI::64::INSTR" - VXI board 0, logical address 64

"VXI1::64::INSTR" - VXI board 1, logical address 64

"GPIB-VXI::64::INSTR" - GPIB-VXI board 0, logical address 64

"GPIB-VXI1::64::INSTR" - GPIB-VXI board 1, logical address 64

"ASRL2::INSTR" - COM port 2

"ChromaIWT" - Logical name "ChromaIWT"

"chr17040E" - Logical Name or Driver Session "chr17040E"

Default Value:  "ASRL1::INSTR"

IDQuery

Variable Type       ViBoolean

Specify whether you want the instrument driver to perform an ID

Query.

Valid Range:

VI_TRUE  (1) - Perform ID Query (Default Value)

VI_FALSE (0) - Skip ID Query

When you set this parameter to VI_TRUE, the driver verifies that the

instrument you initialize is a type that this driver supports.

Circumstances can arise where it is undesirable to send an ID Query

command string to the instrument.  When you set this parameter to

VI_FALSE, the function initializes the instrument without performing

an ID Query.

resetDevice

Variable Type       ViBoolean

Specify whether you want the to reset the instrument during the

initialization procedure.

Valid Range:

VI_TRUE  (1) - Reset Device (Default Value)

VI_FALSE (0) - Don't Reset

instrumentHandle

Variable Type       ViSession (passed by reference)

Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Notes:

(1) This function creates a new session each time you invoke it.

This is useful if you have multiple physical instances of the same

type of instrument.

(2) Avoid creating multiple concurrent sessions to the same physical

instrument.  Although you can create more than one IVI session for

the same resource, it is best not to do so.  A better approach is to

use the same IVI session in multiple execution threads.  You can use

functions chr17040E_LockSession and chr17040E_UnlockSession to

protect sections of code that require exclusive access to the

resource.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_InitiateMeasurement

ViStatus chr17040E_InitiateMeasurement (ViSession instrumentHandle,

ViConstString channelName);

Purpose

This function initiates all measurement and stores result into the driver

cache.

Remote-control command:

MEAS:ALL?

Relative remote-control command:

MEAS:OPER? - Operation Status

MEAS:TIME? - Time ID

MEAS:TEMP? - Temperature Sensor 1 ~ 8

MEAS:VOLT? - Voltage

MEAS:CURR? - Current

MEAS:POW? - Power

MEAS:AH? - Capacity

MEAS:KWH? - Work

MEAS:DCIR? - DCIR

MEAS:ALARM? - Alarm

MEAS:STATe? - Error 1 ~ 3

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_InitWithOptions

ViStatus chr17040E_InitWithOptions (ViRsrc resourceName,

ViBoolean IDQuery,

ViBoolean resetDevice,

ViString optionString,

ViPSession instrumentHandle);

Purpose

This function performs the following initialization actions:

- Creates a new IVI instrument driver and optionally sets the initial

state of the following session attributes:

CHR17040E_ATTR_RANGE_CHECK

CHR17040E_ATTR_QUERY_INSTRUMENT_STATUS

CHR17040E_ATTR_CACHE

CHR17040E_ATTR_SIMULATE

CHR17040E_ATTR_RECORD_COERCIONS

- Opens a session to the specified device using the interface and address

you specify for the Resource Name parameter.

- If the ID Query parameter is set to VI_TRUE, this function queries the

instrument ID and checks that it is valid for this instrument driver.

- If the Reset parameter is set to VI_TRUE, this function resets the

instrument to a known state.

- Sends initialization commands to set the instrument to the state

necessary for the operation of the instrument driver.

- Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Note:  This function creates a new session each time you invoke it.

Although you can open more than one IVI session for the same resource, it

is best not to do so.  You can use the same session in multiple program

threads.  You can use the chr17040E_LockSession and

chr17040E_UnlockSession functions to protect sections of code that

require exclusive access to the resource.

Content from the IVI specifications reproduced with permission from the

IVI Foundation.

The IVI Foundation and its member companies make no warranty of any kind

with regard to this material, including, but not limited to, the implied

warranties of merchantability and fitness for a particular purpose. The

IVI Foundation and its member companies shall not be liable for errors

contained herein or for incidental or consequential damages in connection

with the furnishing, performance, or use of this material.

Parameter List

resourceName

Variable Type       ViRsrc

Pass the resource name of the device to initialize.

You can also pass the name of a driver session or logical name that

you configure with the IVI Configuration utility.  The driver session

identifies a specific device and specifies the initial settings for

the session.  A logical Name identifies a particular driver session.

Refer to the following table below for the exact grammar to use for

this parameter.  Optional fields are shown in square brackets ([]).

Syntax

------------------------------------------------------

GPIB[board]::<primary address>[::secondary address]::INSTR

VXI[board]::<logical address>::INSTR

GPIB-VXI[board]::<logical address>::INSTR

ASRL<port>::INSTR

<LogicalName>

<Driver Session>

If you do not specify a value for an optional field, the following

values are used:

Optional Field - Value

------------------------------------------------------

board - 0

secondary address - none (31)

The following table contains example valid values for this parameter.

"Valid Value" - Description

------------------------------------------------------

"GPIB::22::INSTR" - GPIB board 0, primary address 22 no

secondary address

"GPIB::22::5::INSTR" - GPIB board 0, primary address 22

secondary address 5

"GPIB1::22::5::INSTR" - GPIB board 1, primary address 22

secondary address 5

"VXI::64::INSTR" - VXI board 0, logical address 64

"VXI1::64::INSTR" - VXI board 1, logical address 64

"GPIB-VXI::64::INSTR" - GPIB-VXI board 0, logical address 64

"GPIB-VXI1::64::INSTR" - GPIB-VXI board 1, logical address 64

"ASRL2::INSTR" - COM port 2

"ChromaBatteryPackTestSystem" - Logical name

"ChromaBatteryPackTestSystem"

"chr17040E" - Logical Name or Driver Session "chr17040E"

Default Value:  "TCPIP0::192.168.1.100::2101::SOCKET"

IDQuery

Variable Type       ViBoolean

Specify whether you want the instrument driver to perform an ID

Query.

Valid Range:

VI_TRUE  (1) - Perform ID Query (Default Value)

VI_FALSE (0) - Skip ID Query

When you set this parameter to VI_TRUE, the driver verifies that the

instrument you initialize is a type that this driver supports.

Circumstances can arise where it is undesirable to send an ID Query

command string to the instrument.  When you set this parameter to

VI_FALSE, the function initializes the instrument without performing

an ID Query.

resetDevice

Variable Type       ViBoolean

Specify whether you want the to reset the instrument during the

initialization procedure.

Valid Range:

VI_TRUE  (1) - Reset Device (Default Value)

VI_FALSE (0) - Don't Reset

optionString

Variable Type       ViString

You can use this control to set the initial value of certain

attributes for the session.  The following table lists the attributes

and the name you use in this parameter to identify the attribute.

Name              Attribute Defined Constant

--------------------------------------------

RangeCheck        CHR17040E_ATTR_RANGE_CHECK

QueryInstrStatus  CHR17040E_ATTR_QUERY_INSTRUMENT_STATUS

Cache             CHR17040E_ATTR_CACHE

Simulate          CHR17040E_ATTR_SIMULATE

RecordCoercions   CHR17040E_ATTR_RECORD_COERCIONS

The format of this string is, "AttributeName=Value" where

AttributeName is the name of the attribute and Value is the value to

which the attribute will be set.  To set multiple attributes,

separate their assignments with a comma.

If you pass NULL or an empty string for this parameter, the session

uses the default values for the attributes.   You can override the

default values by assigning a value explicitly in a string you pass

for this parameter.  You do not have to specify all of the attributes

and may leave any of them out.  If you do not specify one of the

attributes, its default value will be used.

The default values for the attributes are shown below:

Attribute Name     Default Value

----------------   -------------

RangeCheck         VI_TRUE

QueryInstrStatus   VI_FALSE

Cache              VI_TRUE

Simulate           VI_FALSE

RecordCoercions    VI_FALSE

The following are the valid values for ViBoolean attributes:

True:     1, True, or VI_TRUE

False:    0, False, or VI_FALSE

Default Value:

"Simulate=0,RangeCheck=1,QueryInstrStatus=0,Cache=1"

Initialize with RS-232 interface:

Append option string with

"DriverSetup=BaudRate:9600;FlowControl:NONE", it is case sensitive.

Valid Values for BaudRate:

9600

19200

38400

57600

115200

Valid Values for FlowControl:

NONE

HARDWARE

Simulate specific instrument model:

Append option string with "DriverSetup=SimulateModel:19501", it is

case sensitive.

Valid Values for SimulateModel:

19501

19501-K

instrumentHandle

Variable Type       ViSession (passed by reference)

Returns a ViSession handle that you use to identify the instrument in

all subsequent instrument driver function calls.

Notes:

(1) This function creates a new session each time you invoke it.

This is useful if you have multiple physical instances of the same

type of instrument.

(2) Avoid creating multiple concurrent sessions to the same physical

instrument.  Although you can create more than one IVI session for

the same resource, it is best not to do so.  A better approach is to

use the same IVI session in multiple execution threads.  You can use

functions chr17040E_LockSession and chr17040E_UnlockSession to

protect sections of code that require exclusive access to the

resource.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_InvalidateAllAttributes

ViStatus chr17040E_InvalidateAllAttributes (ViSession instrumentHandle);

Purpose

This function invalidates the cached values of all attributes for the

session.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_IsNotANumber

ViStatus chr17040E_IsNotANumber (ViSession instrumentHandle,

ViReal64 value,

ViPBoolean isNotANumber);

Purpose

This function returns VI_TRUE if the input value is not a number.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

value

Variable Type       ViReal64

Inputs a value to check if is not a number.

Valid Range:

Double Value

Default Value:

0

isNotANumber

Variable Type       ViBoolean (passed by reference)

Returns VI_TRUE if the input value is not a number.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_LockSession

ViStatus chr17040E_LockSession (ViSession instrumentHandle,

ViPBoolean callerHasLock);

Purpose

This function obtains a multithread lock on the instrument session.

Before it does so, it waits until all other execution threads have

released their locks on the instrument session.

Other threads might have obtained a lock on this session in the following

ways:

- The user's application called chr17040E_LockSession.

- A call to the instrument driver locked the session.

- A call to the IVI engine locked the session.

After your call to chr17040E_LockSession returns successfully, no other

threads can access the instrument session until you call

chr17040E_UnlockSession.

Use chr17040E_LockSession and chr17040E_UnlockSession around a sequence

of calls to instrument driver functions if you require that the

instrument retain its settings through the end of the sequence.

You can safely make nested calls to chr17040E_LockSession within the same

thread.  To completely unlock the session, you must balance each call to

chr17040E_LockSession with a call to chr17040E_UnlockSession.  If,

however, you use the Caller Has Lock parameter in all calls to

chr17040E_LockSession and chr17040E_UnlockSession within a function, the

IVI Library locks the session only once within the function regardless of

the number of calls you make to chr17040E_LockSession.  This allows you

to call chr17040E_UnlockSession just once at the end of the function.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

callerHasLock

Variable Type       ViBoolean (passed by reference)

This parameter serves as a convenience.  If you do not want to use

this parameter, pass VI_NULL.

Use this parameter in complex functions to keep track of whether you

obtain a lock and therefore need to unlock the session.  Pass the

address of a local ViBoolean variable.  In the declaration of the

local variable, initialize it to VI_FALSE.  Pass the address of the

same local variable to any other calls you make to

chr17040E_LockSession or chr17040E_UnlockSession in the same

function.

The parameter is an input/output parameter.  chr17040E_LockSession

and chr17040E_UnlockSession each inspect the current value and take

the following actions:

- If the value is VI_TRUE, chr17040E_LockSession does not lock the

session again.  If the value is VI_FALSE, chr17040E_LockSession

obtains the lock and sets the value of the parameter to VI_TRUE.

- If the value is VI_FALSE, chr17040E_UnlockSession does not attempt

to unlock the session.  If the value is VI_TRUE,

chr17040E_UnlockSession releases the lock and sets the value of the

parameter to VI_FALSE.

Thus, you can, call chr17040E_UnlockSession at the end of your

function without worrying about whether you actually have the lock.

Example:

ViStatus TestFunc (ViSession vi, ViInt32 flags)

{

ViStatus error = VI_SUCCESS;

ViBoolean haveLock = VI_FALSE;

if (flags & BIT_1)

{

viCheckErr( chr17040E_LockSession(vi, &haveLock));

viCheckErr( TakeAction1(vi));

if (flags & BIT_2)

{

viCheckErr( chr17040E_UnlockSession(vi, &haveLock));

viCheckErr( TakeAction2(vi));

viCheckErr( chr17040E_LockSession(vi, &haveLock);

}

if (flags & BIT_3)

viCheckErr( TakeAction3(vi));

}

Error:

/*

At this point, you cannot really be sure that

you have the lock.  Fortunately, the haveLock

variable takes care of that for you.

*/

chr17040E_UnlockSession(vi, &haveLock);

return error;

}

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryBatteryCurrentSpecification

ViStatus chr17040E_QueryBatteryCurrentSpecification

(ViSession instrumentHandle, ViConstString channelName,

ViPReal64 max, ViPReal64 min);

Purpose

This function returns the maximum and minimum of current.

Remote-control command:

BATT:SPEC:ALL?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of current.

Units: A

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of current.

Units: A

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryBatteryResistanceSpecification

ViStatus chr17040E_QueryBatteryResistanceSpecification

(ViSession instrumentHandle, ViConstString channelName,

ViPReal64 max, ViPReal64 min);

Purpose

This function returns the maximum and minimum of internal resistance.

Remote-control command:

BATT:SPEC:ALL?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of internal resistance.

Units: A

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of internal resistance.

Units: A

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryBatteryVoltageProtectionSpecification

ViStatus chr17040E_QueryBatteryVoltageProtectionSpecification

(ViSession instrumentHandle, ViConstString channelName,

ViPReal64 max, ViPReal64 min);

Purpose

This function returns the maximum and minimum of voltage protection

limit.

Remote-control command:

BATT:SPEC:ALL?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of voltage protection limit.

Units: V

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of voltage protection limit.

Units: V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryBatteryVoltageSlewRateSpecification

ViStatus chr17040E_QueryBatteryVoltageSlewRateSpecification

(ViSession instrumentHandle, ViConstString channelName,

ViPReal64 max, ViPReal64 min);

Purpose

This function returns the maximum and minimum of voltage protection

limit.

Remote-control command:

BATT:SPEC:ALL?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of voltage protection limit.

Units: V

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of voltage protection limit.

Units: V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryBatteryVoltageSpecification

ViStatus chr17040E_QueryBatteryVoltageSpecification

(ViSession instrumentHandle, ViConstString channelName,

ViPReal64 max, ViPReal64 min);

Purpose

This function returns the maximum and minimum of voltage.

Remote-control command:

BATT:SPEC:ALL?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of voltage.

Units: V

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of voltage.

Units: V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryCurrentSlewRateSpecification

ViStatus chr17040E_QueryCurrentSlewRateSpecification

(ViSession instrumentHandle, ViConstString channelName,

ViPReal64 max, ViPReal64 min);

Purpose

This function returns the maximum and minimum of current slew rate.

Remote-control command:

SPEC:CURR:SLEW?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of current slew rate.

Units: A/ms

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of current slew rate.

Units: A/ms

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryCurrentSpecification

ViStatus chr17040E_QueryCurrentSpecification (ViSession instrumentHandle,

ViConstString channelName,

ViPReal64 max,

ViPReal64 min);

Purpose

This function returns the maximum and minimum of current.

Remote-control command:

SPEC:CURR?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of current.

Units: A

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of current.

Units: A

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryPowerSpecification

ViStatus chr17040E_QueryPowerSpecification (ViSession instrumentHandle,

ViConstString channelName,

ViPReal64 max,

ViPReal64 min);

Purpose

This function returns the maximum and minimum of power.

Remote-control command:

SPEC:POW?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of power.

Units: W

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of power.

Units: W

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_QueryVoltageSpecification

ViStatus chr17040E_QueryVoltageSpecification (ViSession instrumentHandle,

ViConstString channelName,

ViPReal64 max,

ViPReal64 min);

Purpose

This function returns the maximum and minimum of voltage.

Remote-control command:

SPEC:VOLT:HIGH?

SPEC:VOLT:LOW?

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

max

Variable Type       ViReal64 (passed by reference)

Returns the maximum of voltage.

Units: V

min

Variable Type       ViReal64 (passed by reference)

Returns the minimum of voltage.

Units: V

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ReadInstrData

ViStatus chr17040E_ReadInstrData (ViSession instrumentHandle,

ViInt32 numberOfBytesToRead,

ViChar _VI_FAR readBuffer[],

ViPInt32 numBytesRead);

Purpose

This function reads data from the instrument.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr19311_init or

chr19311_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

numberOfBytesToRead

Variable Type       ViInt32

Pass the maximum number of bytes to read from the instruments.

Valid Range:  0 to the number of elements in the Read Buffer.

Default Value:  256

readBuffer

Variable Type       ViChar[]

After this function executes, this parameter contains the data that

was read from the instrument.

numBytesRead

Variable Type       ViInt32 (passed by reference)

Returns the number of bytes actually read from the instrument and

stored in the Read Buffer.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr19311_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_reset

ViStatus chr17040E_reset (ViSession instrumentHandle);

Purpose

This function resets the instrument to a known state and sends

initialization commands to the instrument.  The initialization commands

set instrument settings such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ResetInterchangeCheck

ViStatus chr17040E_ResetInterchangeCheck (ViSession instrumentHandle);

Purpose

When developing a complex test system that consists of multiple test

modules, it is generally a good idea to design the test modules so that

they can run in any order.  To do so requires ensuring that each test

module completely configures the state of each instrument it uses.  If a

particular test module does not completely configure the state of an

instrument, the state of the instrument depends on the configuration from

a previously executed test module.  If you execute the test modules in a

different order, the behavior of the instrument and therefore the entire

test module is likely to change.  This change in behavior is generally

instrument specific and represents an interchangeability problem.

You can use this function to test for such cases.  After you call this

function, the interchangeability checking algorithms in the specific

driver ignore all previous configuration operations.  By calling this

function at the beginning of a test module, you can determine whether the

test module has dependencies on the operation of previously executed test

modules.

This function does not clear the interchangeability warnings from the

list of previously recorded interchangeability warnings.  If you want to

guarantee that the chr17040E_GetNextInterchangeWarning function only

returns those interchangeability warnings that are generated after

calling this function, you must clear the list of interchangeability

warnings.  You can clear the interchangeability warnings list by

repeatedly calling the chr17040E_GetNextInterchangeWarning function until

no more interchangeability warnings are returned.  If you are not

interested in the content of those warnings, you can call the

chr17040E_ClearInterchangeWarnings function.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ResetTimer

ViStatus chr17040E_ResetTimer (ViSession instrumentHandle,

ViConstString channelName);

Purpose

This function resets the timer.

Remote-control command:

PROCess:TIME:STATe CLEAR

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_ResetWithDefaults

ViStatus chr17040E_ResetWithDefaults (ViSession instrumentHandle);

Purpose

This function resets the instrument and applies initial user specified

settings from the Logical Name which was used to initialize the session.

If the session was created without a Logical Name, this function is

equivalent to the chr17040E_reset function.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_revision_query

ViStatus chr17040E_revision_query (ViSession instrumentHandle,

ViChar _VI_FAR instrumentDriverRevision[],

ViChar _VI_FAR firmwareRevision[]);

Purpose

This function returns the revision numbers of the instrument driver and

instrument firmware.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

instrumentDriverRevision

Variable Type       ViChar[]

Returns the instrument driver software revision numbers in the form

of a string.

You must pass a ViChar array with at least 256 bytes.

firmwareRevision

Variable Type       ViChar[]

Returns the instrument firmware revision numbers in the form of a

string.

You must pass a ViChar array with at least 256 bytes.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_self_test

ViStatus chr17040E_self_test (ViSession instrumentHandle,

ViPInt16 selfTestResult,

ViChar _VI_FAR selfTestMessage[]);

Purpose

This function runs the instrument's self test routine and returns the

test result(s).

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

selfTestResult

Variable Type       ViInt16 (passed by reference)

This control contains the value returned from the instrument self

test.  Zero means success.  For any other code, see the device's

operator's manual.

Self-Test Code    Description

---------------------------------------

0              Passed self test

1              Self test failed

selfTestMessage

Variable Type       ViChar[]

Returns the self-test response string from the instrument. See the

device's operation manual for an explanation of the string's

contents.

You must pass a ViChar array with at least 256 bytes.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetAttributeViBoolean

ViStatus chr17040E_SetAttributeViBoolean (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViBoolean attributeValue);

Purpose

This function sets the value of a ViBoolean attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViBoolean

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViBoolean are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViBoolean

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetAttributeViInt32

ViStatus chr17040E_SetAttributeViInt32 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViInt32 attributeValue);

Purpose

This function sets the value of a ViInt32 attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViInt32 type.

If you choose to see all IVI attributes, the data types appear

to the right of the attribute names in the list box.

Attributes with data types other than ViInt32 are dim.  If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViInt32

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetAttributeViReal64

ViStatus chr17040E_SetAttributeViReal64 (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViReal64 attributeValue);

Purpose

This function sets the value of a ViReal64 attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViReal64

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViReal64 are dim.  If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViReal64

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetAttributeViSession

ViStatus chr17040E_SetAttributeViSession (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViSession attributeValue);

Purpose

This function sets the value of a ViSession attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViSession

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViSession are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViSession

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetAttributeViString

ViStatus chr17040E_SetAttributeViString (ViSession instrumentHandle,

ViChar _VI_FAR channelName[],

ViAttr attributeID,

ViChar _VI_FAR attributeValue[]);

Purpose

This function sets the value of a ViString attribute.

This is a low-level function that you can use to set the values of

instrument-specific attributes and inherent IVI attributes.  If the

attribute represents an instrument state, this function performs

instrument I/O in the following cases:

- State caching is disabled for the entire session or for the particular

attribute.

- State caching is enabled and the currently cached value is invalid or

is different than the value you specify.

This instrument driver contains high-level functions that set most of the

instrument attributes.  It is best to use the high-level driver functions

as much as possible.  They handle order dependencies and multithread

locking for you.  In addition, they perform status checking only after

setting all of the attributes.  In contrast, when you set multiple

attributes using the SetAttribute functions, the functions check the

instrument status after each call.

Also, when state caching is enabled, the high-level functions that

configure multiple attributes perform instrument I/O only for the

attributes whose value you change.  Thus, you can safely call the

high-level functions without the penalty of redundant instrument I/O.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViChar[]

If the attribute is channel-based, this parameter specifies the name

of the channel on which to set the value of the attribute. If the

attribute is not channel-based, then pass VI_NULL or an empty string.

Valid Channel Names:  CH1

Default Value:  ""

attributeID

Variable Type       ViAttr

Pass the ID of an attribute.

From the function panel window, you can use this control as follows.

- Click on the control or press <ENTER>, <spacebar>, or

<ctrl-down arrow>, to display a dialog box containing a

hierarchical list of the available attributes.  Attributes

whose value cannot be set are dim.  Help text is shown for

each attribute.  Select an attribute by double-clicking on it

or by selecting it and then pressing <ENTER>.

Read-only attributes appear dim in the list box.  If you

select a read-only attribute, an error message appears.

A ring control at the top of the dialog box allows you to see

all IVI attributes or only the attributes of the ViString

type.  If you choose to see all IVI attributes, the data types

appear to the right of the attribute names in the list box.

Attributes with data types other than ViString are dim. If

you select an attribute data type that is dim, LabWindows/CVI

transfers you to the function panel for the corresponding

function that is consistent with the data type.

- If you want to enter a variable name, press <CTRL-T> to change

this ring control to a manual input box.

- If the attribute in this ring control has named constants as

valid values, you can view the constants by moving to the

Attribute Value control and pressing <ENTER>.

attributeValue

Variable Type       ViChar[]

Pass the value to which you want to set the attribute.

From the function panel window, you can use this control as follows.

- If the attribute currently showing in the Attribute ID ring

control has constants as valid values, you can view a list of

the constants by pressing <ENTER> on this control.  Select a

value by double-clicking on it or by selecting it and then

pressing <ENTER>.

Note:  Some of the values might not be valid depending on the

current settings of the instrument session.

Default Value: none

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetBatteryOutputState

ViStatus chr17040E_SetBatteryOutputState (ViSession instrumentHandle,

ViConstString channelName,

ViInt32 outputState);

Purpose

This function sets the output state of battery simulator.

Remote-control command:

BATTery:OUTPut

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

outputState

Variable Type       ViInt32

Sets the output state of battery simulator controller.

Valid Values:

CHR17040E_VAL_BATT_OUTPUT_OFF - Off

CHR17040E_VAL_BATT_OUTPUT_MODE_A - Mode A

CHR17040E_VAL_BATT_OUTPUT_MODE_B - Mode B

Default Value: CHR17040E_VAL_BATT_OUTPUT_OFF

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetOutputState

ViStatus chr17040E_SetOutputState (ViSession instrumentHandle,

ViConstString channelName,

ViInt32 outputState);

Purpose

This function sets the output state of charge/discharge controller.

Remote-control command:

OUTPut:STATe

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

outputState

Variable Type       ViInt32

Sets the output state of charge/discharge controller.

Valid Values:

CHR17040E_VAL_OUTPUT_OFF - Off

CHR17040E_VAL_OUTPUT_ON - On

CHR17040E_VAL_OUTPUT_PAUSE - Pause

CHR17040E_VAL_OUTPUT_CONTINUE - Continue

Default Value: CHR17040E_VAL_OUTPUT_OFF

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_SetTimerState

ViStatus chr17040E_SetTimerState (ViSession instrumentHandle,

ViConstString channelName,

ViBoolean timerState);

Purpose

This function sets the timer state.

Remote-control command:

PROCess:TIME:STATe ON|OFF

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

timerState

Variable Type       ViBoolean

Turns on/off the timer.

Valid Values:

VI_TRUE - Turn on timer

VI_FALSE - Turn off timer

Default Value: VI_FALSE

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_UnlockSession

ViStatus chr17040E_UnlockSession (ViSession instrumentHandle,

ViPBoolean callerHasLock);

Purpose

This function releases a lock that you acquired on an instrument session

using chr17040E_LockSession.  Refer to chr17040E_LockSession for

additional information on session locks.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

callerHasLock

Variable Type       ViBoolean (passed by reference)

This parameter serves as a convenience.  If you do not want to use

this parameter, pass VI_NULL.

Use this parameter in complex functions to keep track of whether you

obtain a lock and therefore need to unlock the session.

Pass the address of a local ViBoolean variable.  In the declaration

of the local variable, initialize it to VI_FALSE.  Pass the address

of the same local variable to any other calls you make to

chr17040E_LockSession or chr17040E_UnlockSession in the same

function.

The parameter is an input/output parameter.  chr17040E_LockSession

and chr17040E_UnlockSession each inspect the current value and take

the following actions:

- If the value is VI_TRUE, chr17040E_LockSession does not lock the

session again.  If the value is VI_FALSE, chr17040E_LockSession

obtains the lock and sets the value of the parameter to VI_TRUE.

- If the value is VI_FALSE, chr17040E_UnlockSession does not attempt

to unlock the session.  If the value is VI_TRUE,

chr17040E_UnlockSession releases the lock and sets the value of the

parameter to VI_FALSE.

Thus, you can, call chr17040E_UnlockSession at the end of your

function without worrying about whether you actually have the lock.

Example:

ViStatus TestFunc (ViSession vi, ViInt32 flags)

{

ViStatus error = VI_SUCCESS;

ViBoolean haveLock = VI_FALSE;

if (flags & BIT_1)

{

viCheckErr( chr17040E_LockSession(vi, &haveLock));

viCheckErr( TakeAction1(vi));

if (flags & BIT_2)

{

viCheckErr( chr17040E_UnlockSession(vi, &haveLock));

viCheckErr( TakeAction2(vi));

viCheckErr( chr17040E_LockSession(vi, &haveLock);

}

if (flags & BIT_3)

viCheckErr( TakeAction3(vi));

}

Error:

/*

At this point, you cannot really be sure that

you have the lock.  Fortunately, the haveLock

variable takes care of that for you.

*/

chr17040E_UnlockSession(vi, &haveLock);

return error;

}

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_UploadBatteryCapacityCurve

ViStatus chr17040E_UploadBatteryCapacityCurve

(ViSession instrumentHandle, ViConstString channelName,

ViInt32 numberOfPoints, ViReal64 _VI_FAR capacityCurve[]);

Purpose

This function downloads the capacity curve of battery. All curves must

have the same number of points.

Remote-control command:

BATTery:CURVe

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

numberOfPoints

Variable Type       ViInt32

Inputs the number of curve points.

Valid Values: 1 ~ 150

Default Value: 150

capacityCurve

Variable Type       ViReal64[]

Uploads the capacity curve.

Valid Values: 0 ~ 100%

Default Value: None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_UploadBatteryChargeDCRCurve

ViStatus chr17040E_UploadBatteryChargeDCRCurve

(ViSession instrumentHandle, ViConstString channelName,

ViInt32 numberOfPoints,

ViReal64 _VI_FAR chargeDCRCurve[]);

Purpose

This function uploads the DCR curve when charge of battery. All curves

must have the same number of points.

Remote-control command:

BATTery:CURVe

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

numberOfPoints

Variable Type       ViInt32

Inputs the number of curve points.

Valid Values: 1 ~ 150

Default Value: 150

chargeDCRCurve

Variable Type       ViReal64[]

Downloads the DCR curve when charge. The number of elements must be

equal to the value of Number of Points.

Valid Values: <Spec. Min> ~ <Spec. Max> ohm

Default Value: None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_UploadBatteryDischargeDCRCurve

ViStatus chr17040E_UploadBatteryDischargeDCRCurve

(ViSession instrumentHandle, ViConstString channelName,

ViInt32 numberOfPoints,

ViReal64 _VI_FAR dischargeDCRCurve[]);

Purpose

This function uploads the DCR curve when discharge of battery. All curves

must have the same number of points.

Remote-control command:

BATTery:CURVe

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

numberOfPoints

Variable Type       ViInt32

Inputs the number of curve points.

Valid Values: 1 ~ 150

Default Value: 150

dischargeDCRCurve

Variable Type       ViReal64[]

Uploads the DCR curve when discharge.

Valid Values: <Spec. Min> ~ <Spec. Max> ohm

Default Value: None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_UploadBatteryVoltageCurve

ViStatus chr17040E_UploadBatteryVoltageCurve (ViSession instrumentHandle,

ViConstString channelName,

ViInt32 numberOfPoints,

ViReal64 _VI_FAR voltageCurve[]);

Purpose

This function uploads the voltage curve of battery. All curves must have

the same number of points.

Remote-control command:

BATTery:CURVe

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

channelName

Variable Type       ViConstString

Specifies the channel to set the parameter.

Valid Values:

CH1

CH2

CH3

Default Value: "CH1"

numberOfPoints

Variable Type       ViInt32

Inputs the number of curve points.

Valid Values: 1 ~ 150

Default Value: 150

voltageCurve

Variable Type       ViReal64[]

Uploads the voltage curve.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: None

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_viRead

ViStatus chr17040E_viRead (ViSession instrumentHandle,

ViInt64 numberOfBytesToRead,

ViByte _VI_FAR readBuffer[],

ViPInt64 numberOfBytesRead);

Purpose

This function reads data directly from an instrument using VISA I/O. The

function bypasses the attribute state caching mechanism.

The function assumes that the IVI_ATTR_SYSTEM_IO_SESSION attribute for

the IVI session you specify holds a valid VISA session for the

instrument.

If data is not available at the instrument's output buffer when you call

this function, the instrument might hang up. In that case, the function

does not return until the VISA I/O call times out. If you have disabled

the timeout, the function hangs indefinitely.

Note: This function is supported with the IVI Compliance Package (ICP)

4.6.1 and later. If you are using an earlier version of ICP, you can

still use the WriteInstrData and ReadInstrData functions under the

Obsolete folder.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

numberOfBytesToRead

Variable Type       ViInt64

Specify the maximum number of bytes to read from the instrument. The

Read Buffer parameter must be a ViChar buffer that contains at least

the number of bytes you specify.

If the number of bytes you specify is less than the number of bytes

in the instrument's output buffer, you must call this function again

to empty the output buffer. If you do not empty the instrument's

output buffer, the instrument might return invalid data in response

to subsequent requests.

If data is not available at the instrument's output buffer when you

call this function, the instrument might hang up. In that case, the

function does not return until the VISA I/O call times out. If you

have disabled the timeout, the function hangs indefinitely.

readBuffer

Variable Type       ViByte[]

A buffer in which the function places the data it receives from the

instrument.

The buffer must be a ViChar array that has at least as many bytes as

you specify in the Number Bytes To Read parameter.

This function does not write an ASCII NUL byte to terminate the data,

nor does it clear the buffer beyond the bytes it actually receives

from the instrument.

numberOfBytesRead

Variable Type       ViInt64 (passed by reference)

This control returns the actual number of bytes the function received

from the instrument. This is the value that the VISA viRead function

returns.

If the actual number of bytes received is less than the number of

bytes you specify in the Number Bytes To Read parameter, the

instrument's output buffer has probably emptied.

If the number of bytes received is 0, the most probable cause is that

no data was available at the instrument's output buffer.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_viWrite

ViStatus chr17040E_viWrite (ViSession instrumentHandle,

ViByte _VI_FAR writeBuffer[],

ViInt64 numberOfBytesToWrite,

ViPInt64 numberOfBytesWritten);

Purpose

This function writes a command string directly to an instrument using

VISA I/O. The function bypasses the attribute state caching mechanism and

therefore always invalidates all attribute cache values for the session.

The function assumes that the IVI_ATTR_SYSTEM_IO_SESSION attribute for

the IVI session you specify holds a valid VISA session for the

instrument.

The function calls Ivi_SetNeedToCheckStatus with VI_TRUE.

Note: This function is supported with the IVI Compliance Package (ICP)

4.6.1 and later. If you are using an earlier version of ICP, you can

still use the WriteInstrData and ReadInstrData functions under the

Obsolete folder.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

writeBuffer

Variable Type       ViByte[]

Specify the string you want to send to the instrument.

numberOfBytesToWrite

Variable Type       ViInt64

Specifies the number of bytes to write to the instrument. The value

of this parameter must be less than or equal to the size of buffer.

numberOfBytesWritten

Variable Type       ViInt64 (passed by reference)

Returns the number of bytes actually written.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

chr17040E_WriteInstrData

ViStatus chr17040E_WriteInstrData (ViSession instrumentHandle,

ViChar _VI_FAR writeBuffer[]);

Purpose

This function writes a user-specified string to the instrument.

Note:  This function bypasses IVI attribute state caching.  Therefore,

when you call this function, the cached values for all attributes will be

invalidated.

Parameter List

instrumentHandle

Variable Type       ViSession

The ViSession handle that you obtain from the chr17040E_init or

chr17040E_InitWithOptions function.  The handle identifies a

particular instrument session.

Default Value:  None

writeBuffer

Variable Type       ViChar[]

Pass the string to be written to the instrument.

Return Value

Returns the status code of this operation.  The status code  either

indicates success or describes an error or warning condition.  You

examine the status code from each call to an instrument driver

function to determine if an error occurred.

To obtain a text description of the status code, call the

chr17040E_error_message function.  To obtain additional information

about the error condition, call the chr17040E_GetError function.  To

clear the error information from the driver, call the

chr17040E_ClearError function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

WARNINGS:

None

ERRORS:

None

This instrument driver also returns errors and warnings defined by

other sources.  The following table defines the ranges of additional

status codes that this driver can return.  The table lists the

different include files that contain the defined constants for the

particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFA0000 to 3FFA1FFF     IVI      Warnings

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFA0000 to BFFA1FFF     IVI      Errors

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

Attribute Information for the Following Functions:

chr17040E_SetAttributeViInt32

chr17040E_GetAttributeViInt32

chr17040E_CheckAttributeViInt32

chr17040E_SetAttributeViInt64

chr17040E_GetAttributeViInt64

chr17040E_CheckAttributeViInt64

chr17040E_SetAttributeViReal64

chr17040E_GetAttributeViReal64

chr17040E_CheckAttributeViReal64

chr17040E_SetAttributeViSession

chr17040E_GetAttributeViSession

chr17040E_CheckAttributeViSession

chr17040E_SetAttributeViBoolean

chr17040E_GetAttributeViBoolean

chr17040E_CheckAttributeViBoolean

chr17040E_SetAttributeViString

chr17040E_GetAttributeViString

chr17040E_CheckAttributeViString

Inherent IVI Attributes

User Options

Range Check                              CHR17040E_ATTR_RANGE_CHECK

Query Instrument Status                  CHR17040E_ATTR_QUERY_INSTRUMENT_STATUS

Cache                                    CHR17040E_ATTR_CACHE

Simulate                                 CHR17040E_ATTR_SIMULATE

Record Value Coercions                   CHR17040E_ATTR_RECORD_COERCIONS

Interchange Check                        CHR17040E_ATTR_INTERCHANGE_CHECK

Driver Identification

Description                              CHR17040E_ATTR_SPECIFIC_DRIVER_DESCRIPTION

Driver Prefix                            CHR17040E_ATTR_SPECIFIC_DRIVER_PREFIX

Driver Vendor                            CHR17040E_ATTR_SPECIFIC_DRIVER_VENDOR

Revision                                 CHR17040E_ATTR_SPECIFIC_DRIVER_REVISION

Class Specification Major Version        CHR17040E_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MAJOR_VERSION

Class Specification Minor Version        CHR17040E_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MINOR_VERSION

Driver Capabilities

Supported Instrument Models              CHR17040E_ATTR_SUPPORTED_INSTRUMENT_MODELS

Class Group Capabilities                 CHR17040E_ATTR_GROUP_CAPABILITIES

Instrument Identification

Manufacturer                             CHR17040E_ATTR_INSTRUMENT_MANUFACTURER

Model                                    CHR17040E_ATTR_INSTRUMENT_MODEL

Firmware Revision                        CHR17040E_ATTR_INSTRUMENT_FIRMWARE_REVISION

Advanced Session Information

Logical Name                             CHR17040E_ATTR_LOGICAL_NAME

I/O Resource Descriptor                  CHR17040E_ATTR_IO_RESOURCE_DESCRIPTOR

Driver Setup                             CHR17040E_ATTR_DRIVER_SETUP

Miscellaneous Attributes

ID Query Response                           CHR17040E_ATTR_ID_QUERY_RESPONSE

Configuration

Charge/Discharge Controller

Source Type                              CHR17040E_ATTR_SOURCE_TYPE

Voltage                                  CHR17040E_ATTR_VOLTAGE

Current                                  CHR17040E_ATTR_CURRENT

Power                                    CHR17040E_ATTR_POWER

Cutoff Time                              CHR17040E_ATTR_CUTOFF_TIME

Cutoff Voltage                           CHR17040E_ATTR_CUTOFF_VOLTAGE

Cutoff Current                           CHR17040E_ATTR_CUTOFF_CURRENT

Current Slew Rate                        CHR17040E_ATTR_CURRENT_SLEW_RATE

Output Protection

Over Voltage Protection Enable        CHR17040E_ATTR_OVER_VOLTAGE_PROTECTION_ENABLE

Over Voltage Protection Limit         CHR17040E_ATTR_OVER_VOLTAGE_PROTECTION_LIMIT

Under Voltage Protection Enable       CHR17040E_ATTR_UNDER_VOLTAGE_PROTECTION_ENABLE

Under Voltage Protection Limit        CHR17040E_ATTR_UNDER_VOLTAGE_PROTECTION_LIMIT

Positive Delta Voltage Protection Limit  CHR17040E_ATTR_POSITIVE_DELTA_VOLTAGE_PROTECTION_LIMIT

Negative Delta Voltage Protection Limit  CHR17040E_ATTR_NEGATIVE_DELTA_VOLTAGE_PROTECTION_LIMIT

Over Current Protection Limit         CHR17040E_ATTR_OVER_CURRENT_PROTECTION_LIMIT

Positive Delta Current Protection Limit  CHR17040E_ATTR_POSITIVE_DELTA_CURRENT_PROTECTION_LIMIT

Negative Delta Current Protection Limit  CHR17040E_ATTR_NEGATIVE_DELTA_CURRENT_PROTECTION_LIMIT

Over Capacity Protection Limit        CHR17040E_ATTR_OVER_CAPACITY_PROTECTION_LIMIT

Battery Simulator

Battery Total Capacity                   CHR17040E_ATTR_BATT_TOTAL_CAPACITY

Battery Initial Mode                     CHR17040E_ATTR_BATT_INITIAL_MODE

Battery Parameter Type                   CHR17040E_ATTR_BATT_PARAMETER_TYPE

Battery Initial Capacity                 CHR17040E_ATTR_BATT_INITIAL_CAPACITY

Battery Initial Voltage                  CHR17040E_ATTR_BATT_INITIAL_VOLTAGE

Battery Over Current Protection Limit    CHR17040E_ATTR_BATT_OVER_CURRENT_PROTECTION_LIMIT

Battery Internal Resistance              CHR17040E_ATTR_BATT_INTERNAL_RESISTANCE

Battery Over SOC Warning Limit           CHR17040E_ATTR_BATT_OVER_SOC_WARNING_LIMIT

Battery Under SOC Warning Limit          CHR17040E_ATTR_BATT_UNDER_SOC_WARNING_LIMIT

Battery Over Voltage Warning Limit       CHR17040E_ATTR_BATT_OVER_VOLTAGE_WARNING_LIMIT

Battery Under Voltage Warning Limit      CHR17040E_ATTR_BATT_UNDER_VOLTAGE_WARNING_LIMIT

Battery Full State Voltage               CHR17040E_ATTR_BATT_FULL_STATE_VOLTAGE

Battery Empty State Voltage              CHR17040E_ATTR_BATT_EMPTY_STATE_VOLTAGE

Battery Charge Efficiency                CHR17040E_ATTR_BATT_CHRAGE_EFFICIENCY

Battery Discharge Efficiency             CHR17040E_ATTR_BATT_DISCHRAGE_EFFICIENCY

Battery Over SOC Protection Limit        CHR17040E_ATTR_BATT_OVER_SOC_PROTECTION_LIMIT

Battery Under SOC Protection Limit       CHR17040E_ATTR_BATT_UNDER_SOC_PROTECTION_LIMIT

Battery Over Voltage Protection Limit    CHR17040E_ATTR_BATT_OVER_VOLTAGE_PROTECTION_LIMIT

Battery Under Voltage Protection Limit   CHR17040E_ATTR_BATT_UNDER_VOLTAGE_PROTECTION_LIMIT

Battery Voltage Slew Rate                CHR17040E_ATTR_BATT_VOLTAGE_SLEW_RATE

Action-Status

Output State                                CHR17040E_ATTR_OUTPUT_STATE

Battery Output State                        CHR17040E_ATTR_BATT_OUTPUT_STATE

CHR17040E_ATTR_BATT_CHRAGE_EFFICIENCY

Data Type:     ViReal64

Description:   This attribute specifies the charge efficiency.

Valid Values:  0 ~ 100.0 %

Default Value: 100.0 %

Remote-control command:

BATTery:EFFCHG

CHR17040E_ATTR_BATT_DISCHRAGE_EFFICIENCY

Data Type:     ViReal64

Description:   This attribute specifies the discharge efficiency.

Valid Values:  0 ~ 100.0 %

Default Value: 100.0 %

Remote-control command:

BATTery:EFFDSG

CHR17040E_ATTR_BATT_EMPTY_STATE_VOLTAGE

Data Type:     ViReal64

Description:   This attribute specifies the voltage in empty charge state.

Valid Values:  <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Remote-control command:

BATTery:VL

CHR17040E_ATTR_BATT_FULL_STATE_VOLTAGE

Data Type:     ViReal64

Description:   This attribute specifies the voltage in full charge state.

Valid Values:  <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Remote-control command:

BATTery:VH

CHR17040E_ATTR_BATT_INITIAL_CAPACITY

Data Type:     ViReal64

Description:   This attribute specifies the initial capacity of battery.

Valid Values: 0 ~ 100.0 %

Default Values: 0 %

Remote-control command:

BATTery:INITial:CAP

CHR17040E_ATTR_BATT_INITIAL_MODE

Data Type:     ViInt32

Description:   This attribute specifies the initial mode.

Valid Values:

CHR17040E_VAL_INITIAL_CAPACITY_MODE - Initial Capacity Mode

CHR17040E_VAL_INITIAL_VOLTAGE_MODE - Initial Voltage Mode

Default Value: CHR17040E_VAL_INITIAL_CAPACITY_MODE

Remote-control command:

BATTery:INITial

Values:

CHR17040E_VAL_INITIAL_CAPACITY_MODE            0

Initial capacity mode (SOC or Ah)

CHR17040E_VAL_INITIAL_VOLTAGE_MODE             1

Initial voltage mode

CHR17040E_ATTR_BATT_INITIAL_VOLTAGE

Data Type:     ViReal64

Description:   This attribute specifies the initial voltage.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Remote-control command:

BATTery:INITial:VOL

CHR17040E_ATTR_BATT_INTERNAL_RESISTANCE

Data Type:     ViReal64

Description:   This attribute specifies the internal resistance.

Valid Values: <Spec. Min> ~ <Spec. Max> ohm

Default Value: 1.000 ohm

Remote-control command:

BATTery:ESR

CHR17040E_ATTR_BATT_OUTPUT_STATE

Data Type:     ViInt32

Description:   This attribute specifies the output state of battery simulator.

Valid Values:

CHR17040E_VAL_BATT_OUTPUT_OFF - Stop output

CHR17040E_VAL_BATT_OUTPUT_MODE_A - Start mode A

CHR17040E_VAL_BATT_OUTPUT_MODE_B - Start mode B

Default Value: CHR17040E_VAL_BATT_OUTPUT_OFF

Remote-control command:

BATTery:OUTPut

Values:

CHR17040E_VAL_BATT_OUTPUT_OFF                  0

Stop output

CHR17040E_VAL_BATT_OUTPUT_MODE_A               1

Start mode A

CHR17040E_VAL_BATT_OUTPUT_MODE_B               2

Start mode B

CHR17040E_ATTR_BATT_OVER_CURRENT_PROTECTION_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the over current protection (OCP)

limit.

The battery simulator disables the output when the current is greater than or

equal to the value of this attribute.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Remote-control command:

BATTery:OCP

CHR17040E_ATTR_BATT_OVER_SOC_PROTECTION_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the over SOC protection limit. That is

the BOH parameter of the battery simulator.

The battery simulator disables the output when the SOC is greater than or equal

to the value of this attribute.

Valid Values:  0 ~ 100.0 %

Default Value: 0.0 %

Remote-control command:

BATTery:BOH

CHR17040E_ATTR_BATT_OVER_SOC_WARNING_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the over SOC (state of charge) warning

limit. That is the BCH parameter of the battery simulator.

The battery simulator shows warning when the SOC is greater than or equal to

the value of this attribute.

Valid Values: 0 ~ 100.0 %

Default Value: 100.0 %

Remote-control command:

BATTery:BCH

CHR17040E_ATTR_BATT_OVER_VOLTAGE_PROTECTION_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the over voltage protection limit. That

is the VOH parameter of the battery simulator.

The battery simulator disables the output when the voltage is greater than or

equal to the value of this attribute.

Valid Values:  <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Remote-control command:

BATTery:VOH

CHR17040E_ATTR_BATT_OVER_VOLTAGE_WARNING_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the over voltage warning limit.

The battery simulator shows warning when the voltage is greater than or equal

to the value of this attribute.

Valid Values:  <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Equivalent parameter:

BVH

Remote-control command:

BATTery:BVH

CHR17040E_ATTR_BATT_PARAMETER_TYPE

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   This attribute specifies the parameter type of capacity.

Valid Values:

CHR17040E_VAL_PARAMETER_AH - Ah

CHR17040E_VAL_PARAMETER_SOC - SOC

Default Value: CHR17040E_VAL_PARAMETER_AH

Remote-control command:

BATTery:PARA

Values:

CHR17040E_VAL_PARAMETER_AH                     0

Ah

CHR17040E_VAL_PARAMETER_SOC                    1

SOC

CHR17040E_ATTR_BATT_TOTAL_CAPACITY

Data Type:     ViReal64

Description:   This attribute specifies the total capacity.

Valid Values: 0 ~ 9999.99 Ah

Default Value: 0 Ah

Remote-control command:

BATTery:CAP

CHR17040E_ATTR_BATT_UNDER_SOC_PROTECTION_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the under SOC protection limit. That is

the BOL parameter of the battery simulator.

The battery simulator disables the output when the SOC is less than or equal to

the value of this attribute.

Valid Values:  0 ~ 100.0 %

Default Value: 0 %

Remote-control command:

BATTery:BOL

CHR17040E_ATTR_BATT_UNDER_SOC_WARNING_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the under SOC (state of charge) warning

limit.

The battery simulator shows warning when the SOC is less than or equal to the

value of this attribute.

Valid Values: 0 ~ 100.0 %

Default Value: 0 %

Remote-control command:

BATTery:BCL

CHR17040E_ATTR_BATT_UNDER_VOLTAGE_PROTECTION_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the under voltage protection limit.

That is the VOL parameter of the battery simulator.

The battery simulator disables the output when the voltage is less than or

equal to the value of this attribute.

Valid Values:  <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Remote-control command:

BATTery:VOL

CHR17040E_ATTR_BATT_UNDER_VOLTAGE_WARNING_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the under voltage warning limit.

The battery simulator shows warning when the voltage is less than or equal to

the value of this attribute.

Valid Values:  <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Equivalent parameter:

BVL

Remote-control command:

BATTery:BVL

CHR17040E_ATTR_BATT_VOLTAGE_SLEW_RATE

Data Type:     ViReal64

Description:   This attribute specifies the voltage slew rate of battery

simulator.

Valid Values: <Spec. Min> ~ <Spec. Max> V/ms

Default Value: 170.000 V/ms

Remote-control command:

BATTery:VOLT:SLEW

CHR17040E_ATTR_CACHE

Data Type:     ViBoolean

Description:   Specifies whether to cache the value of attributes.  When

caching is enabled, the instrument driver keeps track of the current instrument

settings and avoids sending redundant commands to the instrument.  Thus, you

can significantly increase execution speed.

The instrument driver can choose always to cache or never to cache

particular attributes regardless of the setting of this attribute.

The default value is VI_TRUE.  Use the chr17040E_InitWithOptions function

to override this value.

CHR17040E_ATTR_CURRENT

Data Type:     ViReal64

Description:   This attribute specifies the output current.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Remote-control command:

SOURce:CURR

CHR17040E_ATTR_CURRENT_SLEW_RATE

Data Type:     ViReal64

Description:   This attribute specifies the current slew rate.

Valid Values: <Spec. Min> ~ <Spec.Max> A/ms

Default Value: 10.00 A/ms

Remote-control command:

SOURce:CURRent:SLEW

CHR17040E_ATTR_CUTOFF_CURRENT

Data Type:     ViReal64

Description:   This attribute specifies the cutoff current.

Valid Values: <Spec. Min> ~ <Spec.Max> A

Default Value: 0 A

Remote-control command:

SOURce:CURRent:CUTOFF

CHR17040E_ATTR_CUTOFF_TIME

Data Type:     ViReal64

Description:   This attribute specifies the cutoff time.

Valid Range:

1 ~ 65535 sec - finite time

CHR17040E_VAL_INFINITE_CUTOFF_TIME - infinite time

Default Value: CHR17040E_VAL_INFINITE_CUTOFF_TIME

Remote-control command:

SOURce:TIME:CUTOFF

CHR17040E_ATTR_CUTOFF_VOLTAGE

Data Type:     ViReal64

Description:   This attribute specifies the cutoff voltage.

Valid Values: <Spec. Min> ~ <Spec.Max> V

Default Value: 0 V

Remote-control command:

SOURce:VOLTage:CUTOFF

CHR17040E_ATTR_DRIVER_SETUP

Data Type:     ViString

Restrictions:  Not settable.

Description:   Some cases exist where you must specify instrument driver

options at initialization time.  An example of this is specifying a particular

instrument model from among a family of instruments that the driver supports.

This is useful when using simulation.  You can specify driver-specific options

through the DriverSetup keyword in the optionsString parameter to the

chr17040E_InitWithOptions function.  If you open an instrument using a logical

name, you can also specify the options through the IVI Configuration Utility.

The default value is an empty string.

CHR17040E_ATTR_GROUP_CAPABILITIES

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains a comma-separated list of class-extension

groups that this driver implements.

CHR17040E_ATTR_ID_QUERY_RESPONSE

Data Type:     ViString

Restrictions:  Not settable.

Description:   Returns the ID Query response string.  The instrument driver

gets the value of this attribute when you pass VI_TRUE for the ID Query

parameter to the chr17040E_init or chr17040E_InitWithOptions function.

CHR17040E_ATTR_INSTRUMENT_FIRMWARE_REVISION

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the firmware revision information for the

instrument you are currently using.

CHR17040E_ATTR_INSTRUMENT_MANUFACTURER

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the name of the instrument manufacturer

you are currently using.

CHR17040E_ATTR_INSTRUMENT_MODEL

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the model number or name of the

instrument that you are currently using.

CHR17040E_ATTR_INTERCHANGE_CHECK

Data Type:     ViBoolean

Description:   Specifies whether to perform interchangeability checking and

retrieve interchangeability warnings.

The default value is VI_FALSE.

Interchangeability warnings indicate that using your application with a

different instrument might cause different behavior.  You call chr17040E Get

Next Interchange Warning to extract interchange warnings.  Call the

chr17040E_ClearInterchangeWarnings function to clear the list of

interchangeability warnings without reading them.

Interchangeability checking logs a warning for each attribute you have not set

that affects the behavior of the instrument.

CHR17040E_ATTR_IO_RESOURCE_DESCRIPTOR

Data Type:     ViString

Restrictions:  Not settable.

Description:   Indicates the resource descriptor the driver uses to identify

the physical device.

If you initialize the driver with a logical name, this attribute contains

the resource descriptor that corresponds to the entry in the IVI Configuration

utility.

If you initialize the instrument driver with the resource descriptor, this

attribute contains that value.

CHR17040E_ATTR_LOGICAL_NAME

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string containing the logical name you specified when opening

the current IVI session.

You may pass a logical name to the chr17040E_init or

chr17040E_InitWithOptions functions.  The IVI Configuration utility must

contain an entry for the logical name.  The logical name entry refers to a

virtual instrument section in the IVI Configuration file.  The virtual

instrument section specifies a physical device and initial user options.

CHR17040E_ATTR_NEGATIVE_DELTA_CURRENT_PROTECTION_LIMIT

Data Type:     ViReal64

Restrictions:  Not gettable.

Description:   This attribute specifies the negative delta current protection

limit.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Equivalent parameter:

UDCP

Remote-control command:

[:SOURce]:CURRent:PROTection:DELTa:UNDer[:LEVel]

CHR17040E_ATTR_NEGATIVE_DELTA_VOLTAGE_PROTECTION_LIMIT

Data Type:     ViReal64

Restrictions:  Not gettable.

Description:   This attribute specifies the negative delta voltage protection

limit.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Equivalent parameter:

UDVP

Remote-control command:

[:SOURce]:VOLTage:PROTection:DELTa:UNDer[:LEVel]

CHR17040E_ATTR_OUTPUT_STATE

Data Type:     ViInt32

Description:   This attribute specifies the output state of charge/discharge

controller.

Valid Values:

CHR17040E_VAL_OUTPUT_OFF - Stop output

CHR17040E_VAL_OUTPUT_ON - Start output

CHR17040E_VAL_OUTPUT_PAUSE - Pause

CHR17040E_VAL_OUTPUT_CONTINUE - Continue

Default Value: CHR17040E_VAL_OUTPUT_OFF

Remote-control command:

OUTPut:STATe

Values:

CHR17040E_VAL_OUTPUT_OFF                       0

Stop output

CHR17040E_VAL_OUTPUT_ON                        1

Start output

CHR17040E_VAL_OUTPUT_PAUSE                     2

Pause

CHR17040E_VAL_OUTPUT_CONTINUE                  3

Continue

CHR17040E_ATTR_OVER_CAPACITY_PROTECTION_LIMIT

Data Type:     ViReal64

Restrictions:  Not gettable.

Description:   This attribute specifies the over capacity protection limit.

The controller disables the output when the capacity is greater than or equal

to the value of this attribute.

Valid Values:  0 ~ <+Inf> Ah

Default Value: 0 Ah

Equivalent parameter:

OQP

Remote-control command:

AH:PROTection

CHR17040E_ATTR_OVER_CURRENT_PROTECTION_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the over current protection (OCP)

limit.

The controller disables the output when the current is greater than or equal to

the value of this attribute.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Remote-control command:

[:SOURce]:CURRent:PROTection[:LEVel]

CHR17040E_ATTR_OVER_VOLTAGE_PROTECTION_ENABLE

Data Type:     ViBoolean

Restrictions:  Not gettable.

Description:   This attribute specifies the enable state of over voltage

protection (OVP).

Valid Values:

VI_TRUE - Enable Protection

VI_FALSE - Disable Protection

Default Value: VI_FALSE

Remote-control command:

[:SOURce]:VOLTage:PROTection:STATe

CHR17040E_ATTR_OVER_VOLTAGE_PROTECTION_LIMIT

Data Type:     ViReal64

Restrictions:  Not gettable.

Description:   This attribute specifies the over voltage protection (OVP)

limit.

The controller disables the output when the voltage is greater than or equal to

the value of this attribute.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Remote-control command:

[:SOURce]:VOLTage:PROTection[:LEVel]

CHR17040E_ATTR_POSITIVE_DELTA_CURRENT_PROTECTION_LIMIT

Data Type:     ViReal64

Description:   This attribute specifies the positive delta current protection

limit.

Valid Values: <Spec. Min> ~ <Spec. Max> A

Default Value: 0 A

Equivalent parameter:

ODCP

Remote-control command:

[:SOURce]:CURRent:PROTection:DELTa[:LEVel]

CHR17040E_ATTR_POSITIVE_DELTA_VOLTAGE_PROTECTION_LIMIT

Data Type:     ViReal64

Restrictions:  Not gettable.

Description:   This attribute specifies the positive delta voltage protection

limit.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Equivalent parameter:

ODVP

Remote-control command:

[:SOURce]:VOLTage:PROTection:DELTa[:LEVel]

CHR17040E_ATTR_POWER

Data Type:     ViReal64

Description:   This attribute specifies the output power.

Valid Value: <Spec. Min> ~ <Spec. Max> W

Default Value: 0 W

Remote-control command:

SOURce:POW

CHR17040E_ATTR_QUERY_INSTRUMENT_STATUS

Data Type:     ViBoolean

Description:   Specifies whether the instrument driver queries the instrument

status after each operation.  Querying the instrument status is very useful for

debugging.  After you validate your program, you can set this attribute to

VI_FALSE to disable status checking and maximize performance

The instrument driver can choose to ignore status checking for particular

attributes regardless of the setting of this attribute.

The default value is VI_FALSE.  Use the chr17040E_InitWithOptions function

to override this value.

CHR17040E_ATTR_RANGE_CHECK

Data Type:     ViBoolean

Description:   Specifies whether to validate attribute values and function

parameters.  If enabled, the instrument driver validates the parameter values

that you pass to driver functions.  Range checking parameters is very useful

for debugging.  After you validate your program, you can set this attribute to

VI_FALSE to disable range checking and maximize performance.

The default value is VI_TRUE.  Use the chr17040E_InitWithOptions function

to override this value.

CHR17040E_ATTR_RECORD_COERCIONS

Data Type:     ViBoolean

Description:   Specifies whether the IVI engine keeps a list of the value

coercions it makes for integer and real type attributes.  You call chr17040E

Get Next Coercion Record to extract and delete the oldest coercion record from

the list.

The default value is VI_FALSE.  Use the chr17040E_InitWithOptions function

to override this value.

CHR17040E_ATTR_SIMULATE

Data Type:     ViBoolean

Description:   Specifies whether or not to simulate instrument driver I/O

operations.  If simulation is enabled, instrument driver functions perform

range checking and call Ivi_GetAttribute and Ivi_SetAttribute functions, but

they do not perform instrument I/O.  For output parameters that represent

instrument data, the instrument driver functions return calculated values.

The default value is VI_FALSE.  Use the chr17040E_InitWithOptions function

to override this value.

CHR17040E_ATTR_SOURCE_TYPE

Data Type:     ViInt32

Description:   This attribute specifies the output source type.

Valid Values:

CHR17040E_VAL_CC_CHARGE - CC Charge

CHR17040E_VAL_CV_CHARGE - CV Charge

CHR17040E_VAL_CP_CHARGE - CP Charge

CHR17040E_VAL_CC_DISCHARGE - CC Discharge

CHR17040E_VAL_CV_DISCHARGE - CV Discharge

CHR17040E_VAL_CP_DISCHARGE - CP Discharge

CHR17040E_VAL_CV_SOURCE - CV Source

CHR17040E_VAL_REST - Rest

Default Value: CHR17040E_VAL_CC_CHARGE

Remote-control command:

SOURce:MODe

Values:

CHR17040E_VAL_CC_CHARGE                        0

Constant Current Charge

CHR17040E_VAL_CV_CHARGE                        1

Constant Voltage Charge

CHR17040E_VAL_CP_CHARGE                        2

Constant Power Charge

CHR17040E_VAL_CC_DISCHARGE                     3

Constant Current Discharge

CHR17040E_VAL_CV_DISCHARGE                     4

Constant Voltage Discharge

CHR17040E_VAL_CP_DISCHARGE                     5

Constant Power Discharge

CHR17040E_VAL_CV_SOURCE                        6

Constant Voltage Source

CHR17040E_VAL_REST                             7

Rest

CHR17040E_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MAJOR_VERSION

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   The major version number of the class specification with which

this driver is compliant.

CHR17040E_ATTR_SPECIFIC_DRIVER_CLASS_SPEC_MINOR_VERSION

Data Type:     ViInt32

Restrictions:  Not settable.

Description:   The minor version number of the class specification with which

this driver is compliant.

CHR17040E_ATTR_SPECIFIC_DRIVER_DESCRIPTION

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains a brief description of the specific

driver.

CHR17040E_ATTR_SPECIFIC_DRIVER_PREFIX

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the prefix for the instrument driver.

The name of each user-callable function in this driver starts with this prefix.

CHR17040E_ATTR_SPECIFIC_DRIVER_REVISION

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains additional version information about this

instrument driver.

CHR17040E_ATTR_SPECIFIC_DRIVER_VENDOR

Data Type:     ViString

Restrictions:  Not settable.

Description:   A string that contains the name of the vendor that supplies this

driver.

CHR17040E_ATTR_SUPPORTED_INSTRUMENT_MODELS

Data Type:     ViString

Restrictions:  Not settable.

Description:   Contains a model code of the instrument. For drivers that

support more than one device, this attribute contains a comma-separated list of

supported instrument models.

CHR17040E_ATTR_UNDER_VOLTAGE_PROTECTION_ENABLE

Data Type:     ViBoolean

Restrictions:  Not gettable.

Description:   This attribute specifies the enable state of under voltage

protection (UVP).

Valid Values:

VI_TRUE - Enable Protection

VI_FALSE - Disable Protection

Default Value: VI_FALSE

Remote-control command:

[:SOURce]:VOLTage:PROTection:UNDer:STATe

CHR17040E_ATTR_UNDER_VOLTAGE_PROTECTION_LIMIT

Data Type:     ViReal64

Restrictions:  Not gettable.

Description:   This attribute specifies the under voltage protection (UVP)

limit.

The controller disables the output when the voltage is less than or equal to

the value of this attribute.

Valid Values: <Spec. Min> ~ <Spec. Max> V

Default Value: 0 V

Remote-control command:

[:SOURce]:VOLTage:PROTection:UNDer[:LEVel]

CHR17040E_ATTR_VOLTAGE

Data Type:     ViReal64

Description:   This attribute specifies the output voltage.

Valid Values: <Spec. Min> ~ <Spec Max.> V

Default Value: 0 V

Remote-control command:

SOURce:VOLT